## Explanation
This file contains information on secondary standard analyses by LAICPMS and EPMA
Reduce to ca - Recovery of BCR on data reduced to Ca by LAICPMS
reduced to Si - as above but for Si
Prefered values - Preferred values of BCR used her

## Reduced to Ca
2020 data
Glasses 28/10/2019 Batch 3 50 um | [9 empty] | MIs 29/10/19 Batch 4 50um | [5 empty] | MIs 29/10/19 Batch 5 50um | [5 empty] | MIs 30/10/19 Batch 2 30um | [5 empty] | MIs 30/10/19 Batch 4 40um | [5 empty] | MIs 30/10/19 Batch 5 25um | [5 empty] | MIs 31/10/19 Batch 5 20um | [5 empty] | >>>>>> | MIs 8/9/20 Batch 2 30um | [5 empty] | MIs 8/9/20 Batch 3 30um | [5 empty] | MIs 8/9/20 Batch 4 40um | [5 empty] | MIs 8/9/20 Batch 5 40um | [7 empty] | MIs 9/9/20 Batch 2 25um | [5 empty] | MIs 9/9/20 Batch 3 25um | [5 empty] | MIs 9/9/20 Batch 4 20um | [5 empty] | MIs 9/9/20 Batch 5 20um | [5 empty] | MIs 11/9/20 Batch 2 50um | [5 empty] | MIs 11/9/20 Batch 2 50um | [5 empty] | MIs 11/9/20 Batch 4 30um | [5 empty] | MIs 11/9/20 Batch 5 25um | [5 empty] | MIs 11/9/20 Batch 6 20um
Reduced to Ca | [9 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [6 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [7 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca | [5 empty] | Reduced to Ca
| BCR_50um_1 | BCR_50um_2 | BCR_50um_3 | BCR_50um_4 | BCR_50um_5 | BCR_50um_6 | BCR_50um_7 | BCR_50um_8 | AVERAGE |  | BCR_50um_1 | BCR_50um_2 | BCR_50um_3 | BCR_50um_4 | AVERAGE |  | BCR_50um_5 | BCR_50um_6 | BCR_50um_7 | BCR_50um_8 | AVERAGE |  | BCR_30um_1 | BCR_30um_2 | BCR_30um_3 | BCR_30um_4 | AVERAGE |  | BCR_40um_1 | BCR_40um_2 | BCR_40um_3 | BCR_40um_4 | AVERAGE |  | BCR_25um_1 | BCR_25um_2 | BCR_25um_3 | BCR_25um_4 | AVERAGE |  | BCR_20um_1 | BCR_20um_2 | BCR_20um_3 | BCR_20um_4 | AVERAGE |  |  | BCR-1 | BCR-2 | BCR-3 | BCR-4 | AVERAGE |  | BCR-5 | BCR-6 | BCR-7 | BCR-8 | AVERAGE |  | BCR-9 | BCR-10 | BCR-11 | BCR-12 | AVERAGE |  | BCR-13 | BCR-14 | BCR-15 | BCR-16 | BCR-17 | BCR-18 | AVERAGE |  | BCR-1 | BCR-2 | BCR-3 | BCR-4 | AVERAGE |  | BCR-5 | BCR-6 | BCR-7 | BCR-8 | AVERAGE |  | BCR-1 | BCR-2 | BCR-3 | BCR-4 | AVERAGE |  | BCR-5 | BCR-6 | BCR-7 | BCR-8 | AVERAGE |  | BCR-1 | BCR-2 | BCR-3 | BCR-4 | AVERAGE |  | BCR-5 | BCR-6 | BCR-7 | BCR-8 | AVERAGE |  | BCR-9 | BCR-10 | BCR-11 | BCR-12 | AVERAGE |  | BCR-13 | BCR-14 | BCR-15 | BCR-16 | AVERAGE |  | BCR-21 | BCR-22 | BCR-23 | BCR-24 | AVERAGE
Li_ppm_m7 | 0.93023 | 1.0625 | 0.98677 | 1.0316 | 1.0561 | 1.0326 | 1.022 | 1.0038 | 1.0157 | Li_ppm_m7 | 0.95584 | 0.9633 | 1.0262 | 1.0113 | 0.98917 | Li_ppm_m7 | 0.96757 | 1.006 | 0.97504 | 0.9761 | 0.98117 | Li_ppm_m7 | 1.0561 | 1.0316 | 1.0444 | 1.0742 | 1.0516 | Li_ppm_m7 | 1.0881 | 1.1095 | 1.0732 | 1.0892 | 1.09 | Na_ppm_m23 | 1.0851 | 1.0934 | 1.06 | 1.0809 | 1.0798 | Mg_ppm_m24 | 3.9481 | 3.9747 | 3.9543 | 3.9903 | 3.9669 |  | Li_ppm_m7 | 1.0518 | 0.96864 | 1.0134 | 1.0732 | 1.0268 | Li_ppm_m7 | 0.9761 | 0.98037 | 1.0209 | 0.96224 | 0.98491 | Li_ppm_m7 | 1.0305 | 0.99637 | 1.0369 | 1.0508 | 1.0286 | Li_ppm_m7 | 1.0316 | 1.0177 | 1.038 | 1.0017 | 1.0636 | 0.9793 | 1.022 | Na_ppm_m23 | 0.95234 | 0.96194 | 0.98907 | 0.95151 | 0.96371 | Na_ppm_m23 | 1.0099 | 1.0304 | 0.98072 | 1.0058 | 1.0067 | Mg_ppm_m24 | 1.1463 | 1.1341 | 1.149 | 1.1803 | 1.1524 | Mg_ppm_m24 | 1.1244 | 1.1383 | 1.123 | 1.1136 | 1.1248 | Li_ppm_m7 | 1.0102 | 0.97077 | 1.0092 | 0.99744 | 0.99691 | Li_ppm_m7 | 1.0412 | 0.98891 | 1.0412 | 0.9793 | 1.0126 | Li_ppm_m7 | 1.1137 | 1.0657 | 1.006 | 0.97717 | 1.0406 | Na_ppm_m23 | 1.0517 | 1.0558 | 1.0225 | 1.035 | 1.0412 | Mg_ppm_m24 | 1.1188 | 1.1411 | 1.1598 | 1.1178 | 1.1344
Be_ppm_m9 | 1.2113 | 0.86314 | 1.4068 | 0.86791 | 1.0443 | 0.93944 | 1.135 | 1.0443 | 1.064 | Be_ppm_m9 | 1.0205 | 1.0348 | 0.76299 | 0.86314 | 0.92036 | Be_ppm_m9 | 1.0825 | 1.0396 | 1.1493 | 0.79161 | 1.0157 | Be_ppm_m9 | 1.0587 | 1.03 | 0.90606 | 1.0396 | 1.0086 | Be_ppm_m9 | 0.97282 | 0.80591 | 0.94897 | 1.0014 | 0.93228 | Mg_ppm_m24 | 1.3694 | 1.2855 | 1.1644 | 1.2576 | 1.2692 | K_ppm_m39 | 43.5823 | 43.4365 | 44.3111 | 43.6552 | 43.7463 |  | Be_ppm_m9 | 0.74392 | 0.84883 | 1.0682 | 0.7773 | 0.85956 | Be_ppm_m9 | 1.1254 | 0.84883 | 0.48641 | 1.4402 | 0.9752 | Be_ppm_m9 | 0.87744 | 0.82022 | 1.0968 | 1.03 | 0.95613 | Be_ppm_m9 | 0.99666 | 0.94421 | 1.0014 | 1.0968 | 1.1111 | 1.1445 | 1.0491 | Mg_ppm_m24 | 1.1206 | 1.0717 | 1.0806 | 1.047 | 1.08 | Mg_ppm_m24 | 1.088 | 1.1034 | 1.1272 | 1.1225 | 1.1103 | Ca_ppm_m43 | 8.4428 | 8.0627 | 8.0899 | 7.9813 | 8.1442 | Ca_ppm_m43 | 1.1195 | 1.0625 | 1.0919 | 1.1268 | 1.1002 | Be_ppm_m9 | 1.2208 | 0.92036 | 0.95851 | 0.87744 | 0.99428 | Be_ppm_m9 | 0.9299 | 0.67239 | 1.2351 | 0.93944 | 0.94421 | Be_ppm_m9 | 0.76299 | 1.1302 | 0.68193 | 0.98236 | 0.88937 | Mg_ppm_m24 | 1.1318 | 1.0573 | 1.0759 | 1.1225 | 1.0969 | Ca_ppm_m43 | 1.1121 | 1.0735 | 1.0441 | 1.1085 | 1.0846
Na_ppm_m23 | 1.0058 | 1.0099 | 1.0892 | 1.0517 | 1.035 | 1.0099 | 1.0266 | 0.97655 | 1.0256 | Na_ppm_m23 | 1.0058 | 1.0542 | 1.035 | 1.0099 | 1.0262 | Na_ppm_m23 | 1.0141 | 1.027 | 1.0517 | 1.0183 | 1.0278 | Na_ppm_m23 | 1.0767 | 1.1059 | 1.0809 | 1.0851 | 1.0871 | Na_ppm_m23 | 0.98072 | 1.0433 | 1.0275 | 1.0433 | 1.0237 | Al_ppm_m27 | 2.0121 | 2.0587 | 1.9797 | 2.0248 | 2.0188 | Ni_ppm_m60 | 1.1338 | 1.1256 | 1.0522 | 1.0685 | 1.095 |  | Na_ppm_m23 | 1.0375 | 1.0275 | 1.0233 | 1.0538 | 1.0355 | Na_ppm_m23 | 1.0333 | 1.0325 | 1.0212 | 1.0195 | 1.0266 | Na_ppm_m23 | 1.0529 | 1.0475 | 0.99032 | 1.017 | 1.0269 | Na_ppm_m23 | 1.0775 | 1.0512 | 1.0108 | 1.027 | 1.035 | 1.0308 | 1.0387 | Al_ppm_m27 | 1.9896 | 1.9952 | 1.9712 | 1.9064 | 1.9656 | Al_ppm_m27 | 1.967 | 1.9769 | 1.9571 | 1.9684 | 1.9674 | Ni_ppm_m60 | 1.0962 | 1.1371 | 1.1453 | 1.2107 | 1.1473 | Ni_ppm_m60 | 0.93259 | 0.98986 | 1.0635 | 0.94077 | 0.98168 | Na_ppm_m23 | 0.99533 | 1.0725 | 1.0325 | 1.0262 | 1.0316 | Na_ppm_m23 | 1.0646 | 1.035 | 1.0204 | 1.06 | 1.045 | Na_ppm_m23 | 1.0642 | 1.0433 | 1.1101 | 1.1101 | 1.0819 | Al_ppm_m27 | 2.0164 | 1.9219 | 1.9092 | 1.9318 | 1.9448 | Ni_ppm_m60 | 1.3007 | 0.95713 | 1.0798 | 0.99804 | 1.0839
Si_ppm_m29 | 1.0147 | 1.0496 | 1.0643 | 1.0147 | 1.1121 | 1.0018 | 1.0368 | 1.0037 | 1.0372 | Si_ppm_m29 | 1.011 | 1.0404 | 1.0809 | 0.99265 | 1.0312 | Si_ppm_m29 | 0.99449 | 1.0368 | 1.1379 | 0.98346 | 1.0381 | Si_ppm_m29 | 0.98346 | 1.1195 | 1.0257 | 1.1158 | 1.0611 | Si_ppm_m29 | 0.96507 | 1.0607 | 1.0625 | 1.0735 | 1.0404 | K_ppm_m39 | 0.98188 | 0.99329 | 1.0141 | 1.0295 | 1.0047 | Sr_ppm_m88 | 1.0047 | 0.97787 | 1.0017 | 1.0077 | 0.998 |  | Si_ppm_m29 | 1.0717 | 0.9614 | 1.0717 | 1.0092 | 1.0285 | Si_ppm_m29 | 1.0331 | 1.0276 | 1.1158 | 1.046 | 1.0556 | Si_ppm_m29 | 1.0938 | 1.0368 | 0.97978 | 1.0349 | 1.0363 | Si_ppm_m29 | 0.99081 | 1.0276 | 1.0221 | 0.99449 | 1.1121 | 1.0625 | 1.0349 | K_ppm_m39 | 0.88322 | 0.92953 | 0.90671 | 0.88456 | 0.90101 | K_ppm_m39 | 0.96443 | 1.0034 | 0.94228 | 0.97852 | 0.97215 | Sr_ppm_m88 | 0.98434 | 0.96084 | 0.96965 | 1.0167 | 0.98288 | Sr_ppm_m88 | 0.96084 | 0.9579 | 0.97553 | 0.96671 | 0.96525 | Si_ppm_m29 | 0.99265 | 1.0625 | 1.0312 | 0.98529 | 1.0179 | Si_ppm_m29 | 1.079 | 1.0423 | 1.0221 | 1.0165 | 1.04 | Si_ppm_m29 | 1.0551 | 1.0147 | 1.1324 | 1.0882 | 1.0726 | K_ppm_m39 | 0.99732 | 0.9698 | 0.98859 | 0.97584 | 0.98289 | Sr_ppm_m88 | 0.94908 | 0.96965 | 0.99316 | 0.95496 | 0.96671
K_ppm_m39 | 0.91745 | 0.94228 | 0.91946 | 0.91409 | 0.89933 | 0.93423 | 0.92819 | 0.89195 | 0.91837 | K_ppm_m39 | 0.97651 | 0.98188 | 0.97718 | 0.96376 | 0.97483 | K_ppm_m39 | 0.82148 | 0.88054 | 0.92013 | 0.95034 | 0.89312 | K_ppm_m39 | 1.002 | 0.97584 | 0.94765 | 1.0161 | 0.9854 | K_ppm_m39 | 0.99463 | 0.97785 | 0.98591 | 1.0013 | 0.98993 | Ti_ppm_m49 | 0.99309 | 1.0002 | 1.0023 | 1.0429 | 1.0096 | Zr_ppm_m90 | 0.97792 | 0.95535 | 0.95874 | 0.96551 | 0.96438 |  | K_ppm_m39 | 1.0087 | 0.97114 | 0.95839 | 1.0208 | 0.98977 | K_ppm_m39 | 0.9651 | 0.98859 | 0.9396 | 0.92013 | 0.95336 | K_ppm_m39 | 1.0544 | 0.99195 | 0.95772 | 0.95839 | 0.9906 | K_ppm_m39 | 0.97584 | 0.97517 | 0.93758 | 0.96913 | 1.0094 | 0.98054 | 0.97461 | Ti_ppm_m49 | 0.95681 | 0.97246 | 1.023 | 0.97388 | 0.98153 | Ti_ppm_m49 | 0.86361 | 0.875 | 0.92622 | 0.92835 | 0.89829 | Zr_ppm_m90 | 1.0029 | 0.99608 | 1.0149 | 1.0789 | 1.0232 | Zr_ppm_m90 | 1.0263 | 0.99493 | 0.99722 | 1.0001 | 1.0046 | K_ppm_m39 | 0.95369 | 0.95101 | 0.98523 | 0.95772 | 0.96191 | K_ppm_m39 | 1.0255 | 1.0174 | 0.99933 | 0.98054 | 1.0057 | K_ppm_m39 | 1.0087 | 1.0268 | 0.94295 | 0.88456 | 0.96577 | Ti_ppm_m49 | 0.95609 | 0.95183 | 0.98739 | 1.0528 | 0.98704 | Zr_ppm_m90 | 1.0132 | 0.97095 | 0.98237 | 0.94125 | 0.97694
Sc_ppm_m45 | 0.96634 | 1.0221 | 1.0545 | 1.0192 | 1.031 | 1.0133 | 1.0133 | 1.0368 | 1.0196 | Sc_ppm_m45 | 0.93403 | 0.98984 | 0.98044 | 0.94842 | 0.96318 | Sc_ppm_m45 | 0.96634 | 0.96634 | 1.0603 | 0.93403 | 0.98176 | Sc_ppm_m45 | 0.92228 | 1.0897 | 0.98396 | 0.9869 | 0.99571 | Sc_ppm_m45 | 0.91053 | 1.0104 | 0.9399 | 0.95165 | 0.95312 | V_ppm_m51 | 1.0764 | 1.0993 | 1.0512 | 1.0764 | 1.0758 | Nb_ppm_m93 | 0.99844 | 0.96639 | 0.9968 | 0.96311 | 0.98119 |  | Sc_ppm_m45 | 0.99571 | 0.96634 | 0.99865 | 0.95753 | 0.97956 | Sc_ppm_m45 | 1.0045 | 0.96634 | 1.0016 | 1.0045 | 0.99424 | Sc_ppm_m45 | 0.95459 | 0.96047 | 0.95165 | 1.0174 | 0.97104 | Sc_ppm_m45 | 0.96282 | 0.99865 | 0.97515 | 0.96781 | 0.99013 | 0.9634 | 0.97633 | V_ppm_m51 | 0.93898 | 0.9344 | 0.97105 | 0.92982 | 0.94356 | V_ppm_m51 | 0.91837 | 0.95502 | 0.93211 | 0.93669 | 0.93555 | Nb_ppm_m93 | 1.0176 | 0.98622 | 0.96889 | 1.0349 | 1.0019 | Nb_ppm_m93 | 0.9986 | 0.97054 | 1.0366 | 0.97962 | 0.99633 | Sc_ppm_m45 | 0.99689 | 1.0145 | 0.99982 | 0.99571 | 1.0017 | Sc_ppm_m45 | 0.99953 | 0.97691 | 0.99571 | 1.0368 | 1.0022 | Sc_ppm_m45 | 0.96928 | 0.93403 | 1.0251 | 0.9869 | 0.97882 | V_ppm_m51 | 0.99166 | 0.98937 | 0.96189 | 0.99853 | 0.98536 | Nb_ppm_m93 | 1.011 | 0.98622 | 0.98952 | 0.97549 | 0.99055
Ti_ppm_m49 | 0.95752 | 0.94542 | 1.0059 | 0.9689 | 1.0272 | 0.94685 | 0.95894 | 0.94258 | 0.96917 | Ti_ppm_m49 | 0.88496 | 0.93262 | 1.0023 | 0.92977 | 0.93742 | Ti_ppm_m49 | 0.57337 | 0.58404 | 0.97246 | 0.87286 | 0.75068 | Ti_ppm_m49 | 0.8629 | 0.99593 | 0.9191 | 0.94969 | 0.93191 | Ti_ppm_m49 | 0.88424 | 1.0607 | 0.94116 | 0.97317 | 0.96481 | Cr_ppm_m53 | 1.2387 | 1.2449 | 1.1204 | 1.164 | 1.192 | Ba_ppm_m137 | 0.9917 | 1.0067 | 1.0007 | 0.98569 | 0.99621 |  | Ti_ppm_m49 | 0.89278 | 0.83089 | 0.92124 | 0.84654 | 0.87286 | Ti_ppm_m49 | 0.91839 | 0.88211 | 0.94898 | 0.90559 | 0.91377 | Ti_ppm_m49 | 0.94329 | 0.94898 | 0.86788 | 0.9376 | 0.92444 | Ti_ppm_m49 | 0.93191 | 0.95396 | 0.96036 | 0.93475 | 0.97744 | 0.95609 | 0.95242 | Cr_ppm_m53 | 0.87146 | 0.74074 | 1.0707 | 0.9835 | 0.91659 | Cr_ppm_m53 | 1.2387 | 0.93993 | 0.93371 | 1.1889 | 1.0753 | Ba_ppm_m137 | 0.99403 | 0.98358 | 0.91493 | 0.96119 | 0.96343 | Ba_ppm_m137 | 0.94776 | 0.93433 | 1 | 1.0104 | 0.97313 | Ti_ppm_m49 | 0.94116 | 0.96961 | 0.95254 | 0.92053 | 0.94596 | Ti_ppm_m49 | 0.97459 | 0.97886 | 0.96819 | 0.97103 | 0.97317 | Ti_ppm_m49 | 0.89634 | 0.85935 | 0.98882 | 0.93618 | 0.92017 | Cr_ppm_m53 | 0.91503 | 1.0582 | 1.0084 | 1.0333 | 1.0037 | Ba_ppm_m137 | 0.95821 | 0.96418 | 0.99701 | 0.98657 | 0.97649
V_ppm_m51 | 0.94815 | 1.01 | 1.01 | 0.94815 | 1.0466 | 0.95502 | 0.97792 | 0.97334 | 0.98364 | V_ppm_m51 | 0.98021 | 0.98708 | 1.0214 | 0.93898 | 0.98193 | V_ppm_m51 | 0.96647 | 1.026 | 1.0329 | 0.93669 | 0.99051 | V_ppm_m51 | 0.91379 | 0.96647 | 0.94585 | 1.0512 | 0.96933 | V_ppm_m51 | 0.98021 | 1.152 | 1.0443 | 1.0718 | 1.0621 | Mn_ppm_m55 | 1.2043 | 1.1902 | 0.99816 | 1.0202 | 1.1032 | La_ppm_m139 | 0.97233 | 0.98654 | 0.9906 | 0.99466 | 0.98603 |  | V_ppm_m51 | 1.0398 | 0.93669 | 0.98479 | 0.94127 | 0.97563 | V_ppm_m51 | 0.95937 | 0.95731 | 1.0627 | 0.98708 | 0.9916 | V_ppm_m51 | 1.0169 | 0.96876 | 0.96876 | 1.0123 | 0.99166 | V_ppm_m51 | 0.95731 | 0.99624 | 1.01 | 0.97792 | 1.0627 | 1.0329 | 1.0062 | Mn_ppm_m55 | 0.9816 | 0.94417 | 0.96871 | 0.9319 | 0.9566 | Mn_ppm_m55 | 0.81779 | 0.85092 | 0.76319 | 0.73558 | 0.79187 | La_ppm_m139 | 0.97871 | 0.96233 | 0.9742 | 0.99918 | 0.9786 | La_ppm_m139 | 0.99918 | 0.97461 | 1.0115 | 0.9828 | 0.99201 | V_ppm_m51 | 0.97792 | 0.98708 | 1.0054 | 0.9596 | 0.9825 | V_ppm_m51 | 1.01 | 1.0169 | 0.99395 | 0.96418 | 0.99624 | V_ppm_m51 | 1.0008 | 0.94356 | 1.0283 | 0.94356 | 0.97906 | Mn_ppm_m55 | 0.96871 | 0.93252 | 1.0632 | 1.1288 | 1.0233 | La_ppm_m139 | 0.96233 | 0.9828 | 0.99509 | 0.9828 | 0.98075
Cr_ppm_m53 | #VALUE! | 1.0582 | 0.94616 | 0.97728 | 0.96483 | 1.0209 | 1.0271 | 0.90881 | 0.98617 | Cr_ppm_m53 | 0.96483 | 0.9835 | 1.0893 | 1.0084 | 1.0115 | Cr_ppm_m53 | 0.92126 | 1.0209 | 0.96483 | 0.95861 | 0.96639 | Cr_ppm_m53 | 0.9835 | 1.1453 | 0.95861 | 1.0955 | 1.0458 | Cr_ppm_m53 | 1.0707 | 1.0333 | 1.1391 | 1.0582 | 1.0753 | Fe_ppm_m56 | 0.88568 | 0.90491 | 1.1656 | 1.2233 | 1.0449 | Ce_ppm_m140 | 0.99218 | 1.0198 | 1.004 | 1.004 | 1.005 |  | Cr_ppm_m53 | 1.2574 | 1.1142 | 0.96483 | 0.97728 | 1.0784 | Cr_ppm_m53 | 0.68472 | 1.0271 | 1.0084 | 1.0458 | 0.94149 | Cr_ppm_m53 | 1.1142 | 0.89013 | 1.0769 | 1.0333 | 1.0286 | Cr_ppm_m53 | 0.91503 | 1.0707 | 1.0707 | 1.0084 | 0.97728 | 1.0644 | 1.0177 | Fe_ppm_m56 | 1.0684 | 1.0032 | 0.87607 | 0.8344 | 0.94551 | Fe_ppm_m56 | 0.90385 | 0.95726 | 0.97009 | 0.93483 | 0.94151 | Ce_ppm_m140 | 1.0044 | 0.9652 | 0.96324 | 0.99266 | 0.98138 | Ce_ppm_m140 | 0.99462 | 0.99462 | 1.0005 | 0.9907 | 0.99512 | Cr_ppm_m53 | 0.9835 | 1.0333 | 1.0644 | 1.0395 | 1.0302 | Cr_ppm_m53 | 1.0458 | 0.98973 | 1.0458 | 0.92748 | 1.0022 | Cr_ppm_m53 | 1.0209 | 0.80921 | 1.0209 | 0.95861 | 0.95238 | Fe_ppm_m56 | 1.1122 | 1.0353 | 1.0929 | 1.11 | 1.0876 | Ce_ppm_m140 | 0.98089 | 0.97305 | 0.99659 | 0.95539 | 0.97648
Mn_ppm_m55 | 0.89693 | 0.95706 | 0.94663 | 0.97178 | 1 | 0.95092 | 0.92638 | 0.95092 | 0.95008 | Mn_ppm_m55 | 0.96503 | 0.96503 | 0.98589 | 0.97791 | 0.97347 | Mn_ppm_m55 | 0.87975 | 0.89693 | 0.94601 | 0.9638 | 0.92163 | Mn_ppm_m55 | 0.96933 | 0.96319 | 0.94233 | 0.95399 | 0.95721 | Mn_ppm_m55 | 0.96933 | 0.98712 | 0.96687 | 0.98589 | 0.9773 | Ni_ppm_m60 | 1.1944 | 1.0471 | 1.1207 | 1.2271 | 1.1473 | Pr_ppm_m141 | 1.0034 | 0.95703 | 0.95858 | 0.98177 | 0.9752 |  | Mn_ppm_m55 | 0.97669 | 0.98344 | 0.94847 | 0.9319 | 0.96012 | Mn_ppm_m55 | 0.92945 | 0.95276 | 0.96748 | 0.94724 | 0.94923 | Mn_ppm_m55 | 0.98466 | 0.94417 | 0.98037 | 0.99877 | 0.97699 | Mn_ppm_m55 | 0.94908 | 0.97914 | 0.96442 | 0.97239 | 1.0098 | 0.99202 | 0.97781 | Ni_ppm_m60 | 0.98168 | 0.98168 | 1.0717 | 1.0389 | 1.0185 | Ni_ppm_m60 | 1.1289 | 1.0389 | 1.088 | 0.98986 | 1.0614 | Pr_ppm_m141 | 0.9624 | 0.96395 | 1.0073 | 0.99335 | 0.98174 | Pr_ppm_m141 | 1.0707 | 0.98561 | 1.0011 | 0.96085 | 1.0046 | Mn_ppm_m55 | 0.94724 | 0.92883 | 0.97546 | 0.94601 | 0.94939 | Mn_ppm_m55 | 0.96564 | 0.95767 | 0.96564 | 0.97423 | 0.9658 | Mn_ppm_m55 | 0.97239 | 0.96994 | 0.95706 | 0.94479 | 0.96104 | Ni_ppm_m60 | 0.97349 | 0.93259 | 1.0144 | 1.0389 | 0.98986 | Pr_ppm_m141 | 1.0104 | 0.95467 | 1.0104 | 0.95931 | 0.98368
Fe_ppm_m56 | 1 | 1.0235 | 1.0684 | 1.0833 | 1.1079 | 1.0385 | 0.99893 | 1.0438 | 1.0455 | Fe_ppm_m56 | 1.0363 | 1.047 | 1.1058 | 1.0833 | 1.0681 | Fe_ppm_m56 | 0.66132 | 0.69551 | 0.94017 | 0.94231 | 0.80983 | Fe_ppm_m56 | 1.0972 | 1.0929 | 1.0502 | 1.015 | 1.0638 | Fe_ppm_m56 | 1.0791 | 1.0353 | 1.0203 | 0.99786 | 1.0331 | Cu_ppm_m65 | 1.1058 | 1.3291 | 1.1657 | 1.0894 | 1.1203 | Nd_ppm_m146 | 0.99093 | 0.95157 | 0.98377 | 0.948 | 0.96857 |  | Fe_ppm_m56 | 1.0513 | 1.0737 | 1.0994 | 1.062 | 1.0716 | Fe_ppm_m56 | 1.0823 | 1.0887 | 1.1528 | 1.1335 | 1.1143 | Fe_ppm_m56 | 1.1036 | 1.063 | 1.1303 | 1.1432 | 1.11 | Fe_ppm_m56 | 1.0812 | 1.1047 | 1.1271 | 1.1036 | 1.1207 | 1.1004 | 1.1063 | Cu_ppm_m65 | 1.1167 | 1.0023 | 0.94237 | 0.87155 | 0.98322 | Cu_ppm_m65 | 1.0677 | 0.9696 | 1.0241 | 1.0295 | 1.0227 | Nd_ppm_m146 | 0.92696 | 0.93778 | 0.98828 | 0.97746 | 0.95762 | Nd_ppm_m146 | 0.95221 | 0.97024 | 0.97746 | 0.97024 | 0.96754 | Fe_ppm_m56 | 1.109 | 1.0705 | 1.0865 | 1.0897 | 1.0889 | Fe_ppm_m56 | 1.0669 | 1.0545 | 1.0652 | 1.1186 | 1.0763 | Fe_ppm_m56 | 1.0983 | 1.0983 | 1.1645 | 1.1645 | 1.1314 | Cu_ppm_m65 | 1.2474 | 1.1603 | 1.1112 | 1.1984 | 1.1793 | Nd_ppm_m146 | 0.91614 | 0.92335 | 0.97024 | 0.91614 | 0.93147
Co_ppm_m59 | 1.0181 | 1.0024 | 1.0942 | 1.0312 | 1.0837 | 1.0076 | 1.0155 | 1.0181 | 1.0338 | Co_ppm_m59 | 0.99974 | 1.0207 | 1.0627 | 1.0076 | 1.0227 | Co_ppm_m59 | 1.0181 | 1.026 | 1.047 | 0.963 | 1.0135 | Co_ppm_m59 | 1.0601 | 1.0758 | 1.0234 | 1.0417 | 1.0502 | Co_ppm_m59 | 1.0102 | 1.0207 | 1.026 | 1.0076 | 1.0161 | Ga_ppm_m71 | 1.119 | 1.0819 | #VALUE! | 1.1933 | 1.1314 | Sm_ppm_m147 | 0.98985 | 0.94059 | 0.96676 | 1.0129 | 0.97753 |  | Co_ppm_m59 | 1.047 | 1.0181 | 1.0811 | 1.0024 | 1.0371 | Co_ppm_m59 | 1.0155 | 0.98137 | 1.1336 | 1.0627 | 1.0483 | Co_ppm_m59 | 1.0234 | 0.99974 | 1.0129 | 1.0732 | 1.0273 | Co_ppm_m59 | 1.0102 | 1.0312 | 1.0575 | 0.99974 | 1.0102 | 1.0338 | 1.0238 | Ga_ppm_m71 | 1.0354 | 1.0215 | 1.054 | 0.97042 | 1.0203 | Ga_ppm_m71 | 0.94721 | 1.0726 | 1.0215 | 1.0865 | 1.0319 | Sm_ppm_m147 | 0.99254 | 0.88692 | 0.94595 | 0.98944 | 0.95371 | Sm_ppm_m147 | 1.0391 | 0.95061 | 1.0252 | 0.9009 | 0.97895 | Co_ppm_m59 | 1.0181 | 1.005 | 1.0391 | 1.0207 | 1.0207 | Co_ppm_m59 | 1.0417 | 1.0129 | 1.0155 | 1.0181 | 1.022 | Co_ppm_m59 | 1.0338 | 1.0024 | 1.1624 | 1.1336 | 1.083 | Ga_ppm_m71 | 1.054 | 1.1144 | 1.0958 | 1.0958 | 1.09 | Sm_ppm_m147 | 0.97701 | 1.1525 | 0.99876 | 0.77664 | 0.97623
Ni_ppm_m60 | #VALUE! | 1.0226 | 1.1535 | 1.0308 | 1.1616 | 0.94895 | 0.93259 | 1.0553 | 1.0436 | Ni_ppm_m60 | 1.0389 | 1.1044 | 1.0062 | 0.92114 | 1.0177 | Ni_ppm_m60 | 0.93995 | 1.0553 | 1.0144 | 0.89987 | 0.97738 | Ni_ppm_m60 | 1.0389 | 1.0226 | 1.0962 | 0.86715 | 1.0062 | Ni_ppm_m60 | 0.98986 | 1.0553 | 1.0144 | 0.97104 | 1.0076 | Ge_ppm_m74 | 1.105 | 0.99448 | 1.1105 | 0.96133 | 1.0428 | Eu_ppm_m153 | 0.98993 | 0.95365 | 0.9692 | 1.0262 | 0.98474 |  | Ni_ppm_m60 | 1.088 | 0.98986 | 1.088 | 1.1371 | 1.0758 | Ni_ppm_m60 | 1.1207 | 1.0717 | 1.0226 | 1.0389 | 1.0635 | Ni_ppm_m60 | 0.9285 | 1.0635 | 0.93259 | 1.1453 | 1.0175 | Ni_ppm_m60 | 0.93177 | 1.0201 | #VALUE! | 1.0103 | 0.99804 | 1.0602 | 1.0041 | Ge_ppm_m74 | 0.97238 | 0.93923 | #VALUE! | 1.3149 | 1.0755 | Ge_ppm_m74 | #VALUE! | 1.0221 | #VALUE! | #VALUE! | 1.0221 | Eu_ppm_m153 | 0.96925 | 0.9432 | 0.82856 | 0.91714 | 0.91454 | Eu_ppm_m153 | 0.97968 | 0.95883 | 1.0474 | 0.97447 | 0.9901 | Ni_ppm_m60 | 1.0553 | 0.99804 | 0.97349 | 0.94404 | 0.99272 | Ni_ppm_m60 | 1.0136 | 0.97349 | 0.95713 | 1.0144 | 0.98965 | Ni_ppm_m60 | 0.98168 | 1.088 | 1.2435 | 1.0717 | 1.0962 | Ge_ppm_m74 | 0.87845 | #VALUE! | 0.79006 | 0.69061 | 0.78637 | Eu_ppm_m153 | 0.88588 | 0.92757 | 1.0839 | 0.98489 | 0.97056
Cu_ppm_m65 | 0.90424 | 0.9805 | 1.0513 | 1.0023 | 1.1712 | 0.90424 | 0.94782 | 0.9805 | 0.99276 | Cu_ppm_m65 | 1.0132 | 0.98377 | 1.0513 | 0.93692 | 0.9963 | Cu_ppm_m65 | 1.0186 | 1.035 | 0.99684 | 0.89334 | 0.98595 | Cu_ppm_m65 | 0.84432 | 0.99684 | 0.93692 | 1.0513 | 0.95735 | Cu_ppm_m65 | 1.0241 | 0.96416 | 1.0241 | 0.99684 | 1.0023 | Rb_ppm_m85 | 0.98149 | 1.0564 | 1.1146 | 1.125 | 1.0693 | Gd_ppm_m157 | 0.94541 | 0.92993 | 0.96862 | 0.88661 | 0.93264 |  | Cu_ppm_m65 | 1.084 | 1.0295 | 1.133 | 1.1276 | 1.0935 | Cu_ppm_m65 | 1.0459 | 0.95871 | 1.0404 | 1.0077 | 1.0132 | Cu_ppm_m65 | 1.0568 | 0.9696 | 1.0077 | 1.1385 | 1.0431 | Cu_ppm_m65 | 0.96743 | 1.1003 | 1.0622 | 0.9805 | 1.0459 | 0.99139 | 1.0246 | Rb_ppm_m85 | 0.96278 | 1.0148 | 0.98773 | 0.94198 | 0.97681 | Rb_ppm_m85 | 0.98357 | 0.98773 | 1.0168 | 1.0771 | 1.0163 | Gd_ppm_m157 | 0.98899 | 0.94497 | 0.87264 | 0.94811 | 0.93868 | Gd_ppm_m157 | 0.95283 | 1.1006 | 1.0346 | 0.99843 | 1.0216 | Cu_ppm_m65 | 0.99139 | 1.0568 | 0.96416 | 0.90424 | 0.97914 | Cu_ppm_m65 | 1.0404 | 0.96416 | 1.0404 | 1.0077 | 1.0132 | Cu_ppm_m65 | 1.1112 | 0.98595 | 1.1494 | 1.0513 | 1.0745 | Rb_ppm_m85 | 1.021 | 1.0501 | 0.94822 | 0.97941 | 0.99969 | Gd_ppm_m157 | 0.95126 | 0.90723 | 1.0189 | 0.91667 | 0.94851
Zn_ppm_m66 | 0.99807 | 1.024 | 0.96566 | 1.024 | 1.0823 | 0.9864 | 0.97214 | 1.0156 | 1.0085 | Zn_ppm_m66 | 1.0233 | 1.0033 | 1.0156 | 0.96566 | 1.002 | Zn_ppm_m66 | 0.94168 | 0.98575 | 1.0214 | 0.94752 | 0.97409 | Zn_ppm_m66 | 1.0175 | 1.037 | 0.95918 | 1.1925 | 1.0515 | Zn_ppm_m66 | 0.99548 | 1.199 | 1.1244 | 1.1219 | 1.1102 | Sr_ppm_m88 | 0.92558 | 0.91089 | 0.96965 | 1.0108 | 0.95423 | Tb_ppm_m159 | 0.92739 | 0.89822 | 0.88514 | 0.98875 | 0.92487 |  | Zn_ppm_m66 | 1.0875 | 1.0253 | 1.0856 | 1.0169 | 1.0538 | Zn_ppm_m66 | 1.0214 | 1.0318 | 1.0629 | 0.97539 | 1.0229 | Zn_ppm_m66 | 0.99288 | 0.97668 | 1.0052 | 0.99936 | 0.99353 | Zn_ppm_m66 | 0.94752 | 1.0233 | 1.0091 | 0.96502 | 0.98381 | 0.92613 | 0.97582 | Sr_ppm_m88 | 0.96995 | 0.93498 | 0.94027 | 0.90001 | 0.9363 | Sr_ppm_m88 | 0.92264 | 0.94615 | 0.95202 | 0.94321 | 0.941 | Tb_ppm_m159 | 1.0071 | 1.0783 | 0.90539 | 0.95626 | 0.98678 | Tb_ppm_m159 | 1.0783 | 1.1089 | 1.1801 | 1.0783 | 1.1114 | Zn_ppm_m66 | 0.91446 | 0.92419 | 0.95853 | 0.8691 | 0.91657 | Zn_ppm_m66 | 0.94363 | 0.9378 | 0.94557 | 0.92678 | 0.93844 | Zn_ppm_m66 | 1.0324 | 0.96502 | 1.0305 | 0.98511 | 1.0033 | Sr_ppm_m88 | 0.99316 | 0.95202 | 0.93116 | 0.94908 | 0.95636 | Tb_ppm_m159 | 0.85453 | 1.0173 | 0.95626 | 0.90539 | 0.93337
Ga_ppm_m71 | 0.92863 | 1.0396 | 1.0261 | 0.989 | 1.0169 | 0.94721 | 0.95185 | 0.99828 | 0.9872 | Ga_ppm_m71 | 1.054 | 1.0122 | 1.041 | 1.0401 | 1.0368 | Ga_ppm_m71 | 1.0169 | 1.0586 | 1.0261 | 0.99364 | 1.0238 | Ga_ppm_m71 | 1.1422 | 1.0726 | 0.989 | 1.054 | 1.0644 | Ga_ppm_m71 | 1.0215 | 1.1561 | 1.0865 | 1.119 | 1.0958 | Y_ppm_m89 | 0.92931 | 0.97224 | 0.96304 | 1.0428 | 0.97684 | Dy_ppm_m163 | 1.0011 | 0.9379 | 0.96706 | 0.99136 | 0.97435 |  | Ga_ppm_m71 | 1.0819 | 0.97971 | 1.0215 | 1.054 | 1.0343 | Ga_ppm_m71 | 0.989 | 0.98435 | 1.0317 | 0.96114 | 0.99155 | Ga_ppm_m71 | 1.0586 | 0.99364 | 1.0261 | 1.0911 | 1.0424 | Ga_ppm_m71 | 1.0257 | 0.98435 | 1.0401 | 1.0169 | 1.0879 | 1.0331 | 1.0313 | Y_ppm_m89 | 0.99065 | 1.0244 | 1.0244 | 0.94464 | 0.99601 | Y_ppm_m89 | 0.96611 | 0.95998 | 1.0152 | 0.99371 | 0.98374 | Dy_ppm_m163 | 1.0202 | 1.0448 | 1.0416 | 1.0777 | 1.0461 | Dy_ppm_m163 | 0.93971 | 1.0169 | 0.99885 | 0.89371 | 0.9623 | Ga_ppm_m71 | 1.0169 | 1.0122 | 1.0958 | 1.0131 | 1.0345 | Ga_ppm_m71 | 1.0354 | 1.0494 | 1.0345 | 1.054 | 1.0433 | Ga_ppm_m71 | 1.0447 | 1.0169 | 1.0215 | 0.989 | 1.018 | Y_ppm_m89 | 1.009 | 0.93544 | 0.9109 | 0.92317 | 0.94464 | Dy_ppm_m163 | 0.89207 | 1.0103 | 1.0366 | 0.91178 | 0.96271
Ge_ppm_m74 | 0.52486 | 0.78453 | 0.86188 | 0.87845 | 0.62983 | 0.74586 | 0.72376 | 0.77348 | 0.74033 | Ge_ppm_m74 | 0.8232 | 0.81768 | 0.82873 | 0.8011 | 0.81768 | Ge_ppm_m74 | 0.85635 | 0.94475 | 0.94475 | 0.83425 | 0.89503 | Ge_ppm_m74 | 1.2873 | 0.70166 | 0.83425 | 0.78453 | 0.90193 | Ge_ppm_m74 | 0.78453 | 0.72928 | 0.91713 | 1.0276 | 0.86464 | Zr_ppm_m90 | 0.9224 | 0.93097 | 0.94867 | 1.0212 | 0.95581 | Ho_ppm_m165 | 0.968 | 0.90504 | 0.89323 | 0.99948 | 0.94144 |  | Ge_ppm_m74 | 0.57459 | 0.8011 | 0.65193 | 1.105 | 0.78315 | Ge_ppm_m74 | 0.53039 | 0.58011 | 0.82873 | 0.74033 | 0.66989 | Ge_ppm_m74 | 0.75691 | 0.86188 | 0.53591 | 0.61326 | 0.69199 | Ge_ppm_m74 | 0.80663 | 0.67956 | 0.52486 | 0.72376 | 0.85635 | 0.93923 | 0.75506 | Zr_ppm_m90 | 0.97951 | 0.97666 | 0.99779 | 0.97323 | 0.9818 | Zr_ppm_m90 | 0.99665 | 1.0126 | 0.97894 | 1.0069 | 0.99879 | Ho_ppm_m165 | 0.94551 | 0.92147 | 1.0096 | 0.97756 | 0.96354 | Ho_ppm_m165 | 1.0657 | 0.99359 | 1.0016 | 0.99359 | 1.0136 | Ge_ppm_m74 | 0.68508 | 0.81768 | 0.85083 | 0.79558 | 0.78729 | Ge_ppm_m74 | 0.64641 | 0.72928 | 0.83978 | 0.64641 | 0.71547 | Ge_ppm_m74 | 0.70166 | 0.67403 | 0.77348 | 0.65746 | 0.70166 | Zr_ppm_m90 | 0.98522 | 0.91269 | 0.9104 | 0.95438 | 0.94068 | Ho_ppm_m165 | 1.0176 | 0.92147 | 0.96955 | 0.97756 | 0.97155
Rb_ppm_m85 | 0.99397 | 1.0023 | 1.0688 | 1.0875 | 1.0064 | 0.97733 | 1.0127 | 0.95654 | 1.0132 | Rb_ppm_m85 | 0.97733 | 1.0189 | 0.98565 | 0.96902 | 0.98773 | Rb_ppm_m85 | 0.98773 | 1.046 | 1.0564 | 0.94822 | 1.0096 | Rb_ppm_m85 | 1.1021 | 1.0938 | 1.1042 | 1.0709 | 1.0927 | Rb_ppm_m85 | 0.97318 | 1.0335 | 1.0148 | 1.0688 | 1.0226 | Nb_ppm_m93 | 1.0605 | 1.058 | 0.92845 | 1.0143 | 1.0153 | Er_ppm_m166 | 0.99527 | 1.0288 | 0.95613 | 0.97011 | 0.98758 |  | Rb_ppm_m85 | 1.048 | 1.0148 | 1.0023 | 1.0584 | 1.0309 | Rb_ppm_m85 | 1.0626 | 1.0397 | 1.0189 | 0.99189 | 1.0283 | Rb_ppm_m85 | 1.0896 | 1.0335 | 0.96694 | 1.0335 | 1.0309 | Rb_ppm_m85 | 1.0376 | 0.99813 | 1.0064 | 0.98149 | 1.0231 | 1.0314 | 1.013 | Nb_ppm_m93 | 0.94165 | 0.97962 | 0.98869 | 0.93175 | 0.96043 | Nb_ppm_m93 | 0.93092 | 0.94413 | 0.96889 | 0.92515 | 0.94227 | Er_ppm_m166 | 1.0045 | 1.0499 | 0.82577 | 0.96765 | 0.96198 | Er_ppm_m166 | 0.9563 | 0.92225 | 1.0414 | 0.93644 | 0.9641 | Rb_ppm_m85 | 1.0085 | 1.0168 | 1.0148 | 1.0335 | 1.0184 | Rb_ppm_m85 | 1.0376 | 1.0252 | 1.0168 | 1.046 | 1.0314 | Rb_ppm_m85 | 1.0543 | 1.0314 | 1.0584 | 1.0085 | 1.0382 | Nb_ppm_m93 | 1.0333 | 1.0077 | 0.9334 | 0.98292 | 0.98931 | Er_ppm_m166 | 0.95914 | 0.95062 | 0.96481 | 0.86266 | 0.93431
Sr_ppm_m88 | 0.96671 | 0.96084 | 1.0314 | 1.0284 | 1.0813 | 0.97259 | 0.97553 | 0.97553 | 0.99904 | Sr_ppm_m88 | 0.93439 | 1.0108 | 0.97553 | 0.91676 | 0.95937 | Sr_ppm_m88 | 0.95026 | 0.94615 | 1.0372 | 0.90207 | 0.95893 | Sr_ppm_m88 | 0.96671 | 1.0255 | 0.95496 | 0.99316 | 0.98508 | Sr_ppm_m88 | 0.91676 | 0.99904 | 0.99904 | 0.9579 | 0.96818 | Mo_ppm_m95 | 1.1817 | 1.1817 | 1.0899 | 1.1258 | 1.1448 | Tm_ppm_m169 | 0.87013 | 0.8459 | 1.0357 | 1.0155 | 0.9418 |  | Sr_ppm_m88 | 0.98464 | 0.92999 | 0.99022 | 0.95702 | 0.96547 | Sr_ppm_m88 | 0.95702 | 0.94615 | 1.0343 | 0.97847 | 0.97898 | Sr_ppm_m88 | 0.96671 | 0.97259 | 0.94027 | 1.0249 | 0.97612 | Sr_ppm_m88 | 0.97759 | 0.98434 | 1.0017 | 0.95408 | 0.99316 | 0.99022 | 0.98351 | Mo_ppm_m95 | 0.94614 | 1.0859 | 1.0499 | 0.99405 | 1.019 | Mo_ppm_m95 | 1.0619 | 1.0699 | 0.96211 | 1.0819 | 1.0439 | Tm_ppm_m169 | 0.97541 | 0.95492 | 0.8873 | 0.96107 | 0.94467 | Tm_ppm_m169 | 0.94262 | 1.1045 | 0.84631 | 0.90369 | 0.94928 | Sr_ppm_m88 | 0.99022 | 1.0131 | 0.99904 | 0.97847 | 0.99522 | Sr_ppm_m88 | 0.99874 | 1.0099 | 0.95202 | 1.0049 | 0.9914 | Sr_ppm_m88 | 0.98728 | 0.96671 | 1.0637 | 1.0666 | 1.0211 | Mo_ppm_m95 | 1.026 | 1.0619 | 1.018 | 1.01 | 1.029 | Tm_ppm_m169 | 0.98361 | 1.0061 | 1.0492 | 0.93443 | 0.99334
Y_ppm_m89 | 0.97531 | 1.0367 | 1.052 | 1.006 | 1.1317 | 0.96918 | 1.052 | 1.0121 | 1.0294 | Y_ppm_m89 | 0.93851 | 0.99065 | 0.99985 | 0.90784 | 0.95921 | Y_ppm_m89 | 0.96918 | 0.95998 | 1.0367 | 0.87103 | 0.95921 | Y_ppm_m89 | 0.8925 | 0.95691 | 0.92624 | 1.0152 | 0.94771 | Y_ppm_m89 | 0.9109 | 1.0397 | 1.006 | 0.97531 | 0.98298 | Sn_ppm_m118 | 1.162 | 1.109 | 0.89682 | 1.1813 | 1.0873 | Yb_ppm_m172 | 0.90716 | 0.88048 | 1.0139 | 0.91013 | 0.92792 |  | Y_ppm_m89 | 1.0305 | 0.98144 | 1.052 | 0.99678 | 1.0152 | Y_ppm_m89 | 1.0182 | 0.96918 | 1.0397 | 0.99371 | 1.0052 | Y_ppm_m89 | 0.98144 | 0.98144 | 0.94771 | 1.037 | 0.98689 | Y_ppm_m89 | 0.95936 | 0.98758 | 0.99065 | 0.94771 | 1.0121 | 0.99065 | 0.98134 | Sn_ppm_m118 | 1.0608 | 1.0415 | 1.0029 | 0.892 | 0.99928 | Sn_ppm_m118 | 1.08 | 1.1427 | 1.1041 | 1.0993 | 1.1066 | Yb_ppm_m172 | 0.85637 | 0.90144 | 0.86538 | 1.0397 | 0.91572 | Yb_ppm_m172 | 0.94952 | 1.0036 | 1.1178 | 0.97055 | 1.0104 | Y_ppm_m89 | 1.0244 | 1.0244 | 1.0014 | 0.99678 | 1.0117 | Y_ppm_m89 | 1.0581 | 1.0305 | 1.0213 | 1.0274 | 1.0344 | Y_ppm_m89 | 0.98758 | 0.92624 | 1.0612 | 1.0428 | 1.0044 | Sn_ppm_m118 | 0.97878 | 1.1524 | 1.1041 | 1.0463 | 1.0704 | Yb_ppm_m172 | 0.98257 | 0.9375 | 0.92548 | 1.0156 | 0.96529
Zr_ppm_m90 | 0.97666 | 1.0223 | 1.0109 | 0.97209 | 1.0738 | 0.9698 | 1.0166 | 1.0006 | 1.0054 | Zr_ppm_m90 | 0.93611 | 0.98066 | 0.99665 | 0.92525 | 0.95967 | Zr_ppm_m90 | 0.95153 | 0.98237 | 1.0332 | 0.89099 | 0.96452 | Zr_ppm_m90 | 0.90241 | 0.96923 | 0.91269 | 1.0115 | 0.94896 | Zr_ppm_m90 | 0.92525 | 1.108 | 0.97095 | 0.95153 | 0.98894 | Ba_ppm_m137 | 0.90746 | 0.90896 | 0.98358 | 1.0836 | 0.9709 | Lu_ppm_m175 | 0.94339 | 0.86988 | 0.9924 | 0.98832 | 0.9485 |  | Zr_ppm_m90 | 1.0355 | 0.93439 | 1.0229 | 0.97437 | 0.99179 | Zr_ppm_m90 | 0.99094 | 0.94125 | 1.0195 | 0.99665 | 0.98708 | Zr_ppm_m90 | 0.98408 | 0.96009 | 0.94296 | 1.0212 | 0.97709 | Zr_ppm_m90 | 0.95381 | 0.98751 | 0.98294 | 0.96409 | 1.0481 | 1.0064 | 0.99046 | Ba_ppm_m137 | 0.94179 | 0.93134 | 0.96119 | 0.94478 | 0.94478 | Ba_ppm_m137 | 0.95224 | 0.98209 | 0.98507 | 0.96269 | 0.97052 | Lu_ppm_m175 | 0.93814 | 1.1031 | 0.86804 | 0.93814 | 0.96186 | Lu_ppm_m175 | 1.1072 | 1.0825 | 0.97526 | 1.0186 | 1.0459 | Zr_ppm_m90 | 1.0104 | 1.0126 | 1.0149 | 0.98237 | 1.0051 | Zr_ppm_m90 | 1.0452 | 0.98351 | 0.9995 | 0.9818 | 1.0025 | Zr_ppm_m90 | 0.98066 | 0.94296 | 1.0452 | 0.96638 | 0.9838 | Ba_ppm_m137 | 0.98358 | 0.94478 | 0.98507 | 1.0209 | 0.98358 | Lu_ppm_m175 | 1.0392 | 0.9299 | 0.93608 | 0.89691 | 0.95052
Nb_ppm_m93 | 0.95156 | 1.0052 | 0.93918 | 0.95321 | 1.0225 | 1.0019 | 0.96806 | 0.96476 | 0.9758 | Nb_ppm_m93 | 0.96971 | 0.99199 | 0.9953 | 0.94825 | 0.97631 | Nb_ppm_m93 | 0.96063 | 1.0126 | 1.0102 | 0.97219 | 0.9889 | Nb_ppm_m93 | 0.95733 | 0.96146 | 0.97219 | 1.0267 | 0.97941 | Nb_ppm_m93 | 0.94908 | 1.0894 | 1.0135 | 1.0407 | 1.0231 | La_ppm_m139 | 0.98649 | 0.95414 | 0.96233 | 1.0401 | 0.98577 | [7 empty] | Nb_ppm_m93 | 1.0176 | 0.92185 | 0.97054 | 0.98292 | 0.97322 | Nb_ppm_m93 | 0.95898 | 0.97054 | 1.0159 | 0.94413 | 0.97239 | Nb_ppm_m93 | 1.0258 | 0.96228 | 0.95238 | 1.0143 | 0.98869 | Nb_ppm_m93 | 0.94 | 0.98374 | 0.97796 | 0.96724 | 1.0481 | 0.9986 | 0.98594 | La_ppm_m139 | 1.0156 | 0.9828 | 0.99509 | 0.93776 | 0.9828 | La_ppm_m139 | 0.94185 | 0.94185 | 0.94595 | 0.96642 | 0.94902 | [12 empty] | Nb_ppm_m93 | 0.97962 | 0.96806 | 0.99695 | 0.95816 | 0.9757 | Nb_ppm_m93 | 1.0044 | 0.99034 | 0.99612 | 0.92185 | 0.97817 | Nb_ppm_m93 | 1.0019 | 0.96063 | 0.96559 | 0.89791 | 0.95651 | La_ppm_m139 | 1.0238 | 0.95823 | 0.90909 | 0.97052 | 0.9654
[40 empty] | Th/U | #REF! | #REF! | #REF! | #REF! | #REF!
Mo_ppm_m95 | 1.0499 | 1.0539 | 1.0699 | 1.042 | 0.92219 | 1.0539 | 1.018 | 1.018 | 1.0285 | Mo_ppm_m95 | 1.0567 | 1.0799 | 1.0595 | 1.0459 | 1.0605 | Mo_ppm_m95 | 1.0579 | 1.0939 | 1.0619 | 1.0891 | 1.0757 | Mo_ppm_m95 | 1.1338 | 1.1857 | 1.1537 | 1.1218 | 1.1487 | Mo_ppm_m95 | 1.0978 | 1.0659 | 1.0779 | 1.0899 | 1.0829 | Ce_ppm_m140 | 1.0241 | 0.99266 | 1.0103 | 1.0751 | 1.0255 | [7 empty] | Mo_ppm_m95 | 1.034 | 0.99924 | 0.96091 | 1.0683 | 1.0156 | Mo_ppm_m95 | 1.1222 | 1.0859 | 0.95652 | 0.92538 | 1.0225 | Mo_ppm_m95 | 1.0739 | 1.0579 | 0.93097 | 0.98047 | 1.0108 | Mo_ppm_m95 | 1.0112 | 0.99006 | 0.89664 | 1.0559 | 1.0899 | 1.0899 | 1.0223 | Ce_ppm_m140 | 0.98874 | 0.98482 | 0.97108 | 0.91615 | 0.9652 | Ce_ppm_m140 | 0.924 | 0.93381 | 0.95343 | 0.97893 | 0.94754 | [12 empty] | Mo_ppm_m95 | 1.0172 | 1.0699 | 1.0711 | 1.0579 | 1.054 | Mo_ppm_m95 | 1.0978 | 1.0899 | 1.0751 | 1.0739 | 1.0842 | Mo_ppm_m95 | 1.042 | 1.0539 | 1.0699 | 0.99804 | 1.041 | Ce_ppm_m140 | 0.9907 | 0.94754 | 0.96324 | 0.95147 | 0.96324
Cd_ppm_m111 | #VALUE! | 0.61446 | #VALUE! | 0.60463 | 0.44732 | 0.69802 | 0.36867 | 1.0323 | 0.62757 | Cd_ppm_m111 | 0.44732 | 0.88482 | 0.73735 | 1.2289 | 0.8246 | Cd_ppm_m111 | #VALUE! | 1.2289 | 1.5239 | 1.2289 | 1.3272 | Cd_ppm_m111 | 0.68819 | #VALUE! | #VALUE! | 1.6222 | 1.1552 | Cd_ppm_m111 | 1.8679 | 1.0323 | #VALUE! | #VALUE! | 1.4501 | Pr_ppm_m141 | 0.98097 | 0.95002 | 0.99954 | 1.0521 | 0.99567 | [7 empty] | Cd_ppm_m111 | #VALUE! | 0.2212 | 0.12289 | #VALUE! | 0.17205 | Cd_ppm_m111 | 0.93397 | #VALUE! | #VALUE! | #VALUE! | 0.93397 | Cd_ppm_m111 | #VALUE! | 0.58496 | 0.73735 | #VALUE! | 0.66116 | Cd_ppm_m111 | 0.16713 | 0.74226 | #VALUE! | 0.67836 | 0.89465 | 0.70294 | 0.63707 | Pr_ppm_m141 | 0.97787 | 0.9918 | 0.9918 | 0.92372 | 0.9713 | Pr_ppm_m141 | 0.94383 | 0.95157 | 0.9624 | 1.0135 | 0.96782 | [12 empty] | Cd_ppm_m111 | 0.67344 | 0.78159 | 0.93397 | #VALUE! | 0.79634 | Cd_ppm_m111 | #VALUE! | #VALUE! | 1.0323 | 1.3272 | 1.1798 | Cd_ppm_m111 | #VALUE! | 0.68819 | 0.93397 | 0.49157 | 0.70458 | Pr_ppm_m141 | 0.95776 | 0.95312 | 0.90979 | 0.94693 | 0.9419
In_ppm_m115 | 1.0506 | 1.122 | 1.173 | 0.96896 | 3.1619 | 0.85677 | 1.0914 | 1.0302 | 1.3068 | In_ppm_m115 | 1.173 | 1.071 | 1.2647 | 0.99956 | 1.1271 | In_ppm_m115 | 1.1526 | 1.1934 | 1.2036 | 0.97916 | 1.1322 | In_ppm_m115 | 1.2036 | 1.2138 | 1.122 | 0.79557 | 1.0837 | In_ppm_m115 | 0.87716 | 1.2342 | 1.1934 | 1.2953 | 1.15 | Nd_ppm_m146 | 0.95221 | 0.91975 | 0.93417 | 1.0063 | 0.95311 | [7 empty] | In_ppm_m115 | 2.1419 | 2.5091 | 1.9685 | 1.5605 | 2.045 | In_ppm_m115 | 0.95876 | 0.91796 | 2.4479 | 2.4479 | 1.6931 | In_ppm_m115 | 1.8359 | 1.6523 | 1.9787 | 2.4581 | 1.9813 | In_ppm_m115 | 1.8461 | 1.9889 | 2.8049 | 0.99956 | 1.0404 | 1.0098 | 1.6149 | Nd_ppm_m146 | 0.95221 | 0.99188 | 0.9486 | 0.93417 | 0.95672 | Nd_ppm_m146 | 0.9486 | 1.0027 | 1.0135 | 1.0135 | 0.99459 | [12 empty] | In_ppm_m115 | 1.4687 | 0.92816 | 1.173 | 1.122 | 1.173 | In_ppm_m115 | 0.87716 | 0.83637 | 1.1016 | 1.6625 | 1.1194 | In_ppm_m115 | 1.9277 | 1.1424 | 2.0603 | 2.3765 | 1.8767 | Nd_ppm_m146 | 1.0099 | 0.99188 | 0.90532 | 0.9991 | 0.97656
Sn_ppm_m118 | 1.0029 | 0.98361 | 0.99807 | 0.99807 | 1.0752 | 0.95468 | 0.97878 | 0.92093 | 0.98903 | Sn_ppm_m118 | 1.0077 | 1.027 | 1.0318 | 0.99807 | 1.0162 | Sn_ppm_m118 | 1.0897 | 1.08 | 0.85824 | 0.97396 | 1.0005 | Sn_ppm_m118 | 1.0752 | 1.0415 | 0.99807 | 0.98361 | 1.0246 | Sn_ppm_m118 | 1.1716 | 1.215 | 1.2777 | 1.1186 | 1.1958 | Sm_ppm_m147 | 0.89779 | 0.92731 | 0.758 | 1.0391 | 0.90556 | [7 empty] | Sn_ppm_m118 | 1.1813 | 1.1475 | 0.97396 | 1.0849 | 1.0969 | Sn_ppm_m118 | 1.0366 | 1.1234 | 1.1186 | 1.0897 | 1.0921 | Sn_ppm_m118 | 0.92093 | 0.88717 | 1.0222 | 1.0752 | 0.97637 | Sn_ppm_m118 | 1.0318 | 1.0511 | 1.0415 | 1.0656 | 1.1041 | 1.0897 | 1.064 | Sm_ppm_m147 | 0.98944 | 1.0236 | 1.0454 | 0.82479 | 0.9708 | Sm_ppm_m147 | 1.05 | 0.9739 | 1.0019 | 1.0547 | 1.0201 | [12 empty] | Sn_ppm_m118 | 0.98361 | 0.99807 | 1.027 | 0.99807 | 1.0017 | Sn_ppm_m118 | 0.99325 | 1.0849 | 1.0608 | 0.92093 | 1.0149 | Sn_ppm_m118 | 0.99807 | 0.93057 | 1.0029 | 0.95468 | 0.97155 | Sm_ppm_m147 | 1.0795 | 0.99876 | 0.95216 | 1.1153 | 1.0364
Sb_ppm_m121 | 1.0312 | 1.0935 | 1.0218 | 0.93458 | 1.2274 | 0.89097 | 0.76636 | 0.97508 | 0.9926 | Sb_ppm_m121 | 0.92212 | 1.081 | 1.1651 | 1.1745 | 1.0857 | Sb_ppm_m121 | 0.93769 | 0.87539 | 1.2274 | 0.82866 | 0.96729 | Sb_ppm_m121 | 0.90966 | 1.3396 | 0.90343 | 0.98131 | 1.0335 | Sb_ppm_m121 | 1.2305 | 0.86604 | 1.4953 | 0.98131 | 1.1433 | Eu_ppm_m153 | 0.91714 | 0.89109 | 0.96925 | 1.0578 | 0.95883 | [7 empty] | Sb_ppm_m121 | 1.2118 | 0.80062 | 0.94704 | 0.96885 | 0.98209 | Sb_ppm_m121 | 1.0592 | 0.97196 | 1.0841 | 0.98131 | 1.0241 | Sb_ppm_m121 | 0.78816 | 0.95639 | 0.94393 | 1.134 | 0.95561 | Sb_ppm_m121 | 0.99065 | 0.89097 | 0.92835 | 0.90654 | 0.83801 | 1.028 | 0.93043 | Eu_ppm_m153 | 0.97968 | 1.0266 | 0.93278 | 0.88067 | 0.95492 | Eu_ppm_m153 | 1.0787 | 0.93278 | 0.97968 | 0.96925 | 0.9901 | [12 empty] | Sb_ppm_m121 | 0.96885 | 0.86293 | 0.90031 | 0.83489 | 0.89174 | Sb_ppm_m121 | 0.80062 | 0.99065 | 0.97508 | 1.0031 | 0.94237 | Sb_ppm_m121 | 0.86916 | 1.0779 | 0.89097 | 1.3364 | 1.0436 | Eu_ppm_m153 | 0.9901 | 0.94841 | 0.83898 | 0.89109 | 0.91714
Cs_ppm_m133 | 0.93299 | 1.0602 | 0.97201 | 0.95929 | 1.0178 | 0.91688 | 1.0178 | 1.0136 | 0.98632 | Cs_ppm_m133 | 0.98473 | 1.011 | 1.0534 | 0.97625 | 1.0064 | Cs_ppm_m133 | 0.93469 | 1 | 1.0789 | 0.92536 | 0.98473 | Cs_ppm_m133 | 1.0178 | 1.0517 | 1.0178 | 1.145 | 1.0581 | Cs_ppm_m133 | 1.0517 | 1.162 | 1.0254 | 1.0602 | 1.0749 | Gd_ppm_m157 | 1.0016 | 0.8978 | 0.95126 | 0.94811 | 0.94969 | [7 empty] | Cs_ppm_m133 | 1.0271 | 1.0178 | 0.99237 | 1.0857 | 1.0307 | Cs_ppm_m133 | 1.0331 | 0.9788 | 1.0772 | 1.0348 | 1.031 | Cs_ppm_m133 | 1.1137 | 1.0076 | 0.95165 | 0.97455 | 1.0119 | Cs_ppm_m133 | 1.0136 | 1.011 | 1.0297 | 0.97371 | 1.0746 | 1.0399 | 1.0237 | Gd_ppm_m157 | 1.0252 | 1.0016 | 1.0975 | 1.022 | 1.0366 | Gd_ppm_m157 | 0.92138 | 0.97327 | 0.90094 | 0.98585 | 0.94536 | [12 empty] | Cs_ppm_m133 | 1.0093 | 1.0127 | 0.99915 | 0.96353 | 0.99618 | Cs_ppm_m133 | 1.0187 | 1.0263 | 1.0297 | 0.95589 | 1.0076 | Cs_ppm_m133 | 1.0136 | 0.93808 | 0.98388 | 0.94572 | 0.97031 | Gd_ppm_m157 | 1.0393 | 0.96855 | 0.96541 | 0.92767 | 0.97524
Ba_ppm_m137 | 0.97761 | 1.0075 | 1.0134 | 1.0388 | 0.98955 | 0.98507 | 0.98358 | 1.0179 | 1.0017 | Ba_ppm_m137 | 1.003 | 1.0045 | 1.0075 | 1.0015 | 1.0041 | Ba_ppm_m137 | 0.95821 | 0.96866 | 0.99403 | 0.95224 | 0.96828 | Ba_ppm_m137 | 0.98657 | 0.99552 | 0.99104 | 0.98955 | 0.99067 | Ba_ppm_m137 | 1.0269 | 1.0239 | 0.97612 | 0.99254 | 1.0049 | Tb_ppm_m159 | 0.9766 | 0.9766 | 0.92574 | 1.0478 | 0.98169 | [7 empty] | Ba_ppm_m137 | 0.9597 | 0.96269 | 0.96418 | 1.006 | 0.97313 | Ba_ppm_m137 | 1.0119 | 0.98358 | 1.0015 | 0.98358 | 0.99515 | Ba_ppm_m137 | 0.99552 | 0.98955 | 0.9806 | 1 | 0.99142 | Ba_ppm_m137 | 1.003 | 0.99701 | 0.97761 | 0.98657 | 0.99701 | 0.98657 | 0.99129 | Tb_ppm_m159 | 1.0071 | 0.87182 | 1.0173 | 0.93591 | 0.95804 | Tb_ppm_m159 | 0.83418 | 0.84435 | 1.0173 | 0.92574 | 0.90539 | [12 empty] | Ba_ppm_m137 | 0.99254 | 1.0015 | 1.0313 | 1.0328 | 1.0146 | Ba_ppm_m137 | 1.0328 | 1.0448 | 1.0194 | 1.0358 | 1.0332 | Ba_ppm_m137 | 0.95672 | 0.98806 | 0.97612 | 0.98358 | 0.97612 | Tb_ppm_m159 | 0.9766 | 0.99695 | 0.86673 | 0.94608 | 0.94659
La_ppm_m139 | 0.9869 | 0.9869 | 1.0238 | 1.0278 | 1.0811 | 0.99509 | 0.98812 | 1.0156 | 1.0132 | La_ppm_m139 | 0.97052 | 1.0127 | 1.0012 | 0.93079 | 0.97881 | La_ppm_m139 | 0.96233 | 0.95414 | 1.0115 | 0.92547 | 0.96335 | La_ppm_m139 | 0.91728 | 1.0565 | 1.0074 | 1.0278 | 1.0023 | La_ppm_m139 | 0.92138 | 1.0524 | 0.97871 | 0.95536 | 0.97697 | Dy_ppm_m163 | 0.93149 | 0.93971 | 0.95285 | 1.0333 | 0.96435 | [7 empty] | La_ppm_m139 | 0.94103 | 0.94758 | 0.96355 | 0.95414 | 0.95158 | La_ppm_m139 | 1.027 | 0.99713 | 1.0082 | 0.99058 | 1.0057 | La_ppm_m139 | 0.98853 | 0.9869 | 0.94349 | 1.0037 | 0.98065 | La_ppm_m139 | 0.96929 | 0.99386 | 1.0041 | 0.97093 | 0.9869 | 0.97011 | 0.98253 | Dy_ppm_m163 | 0.97749 | 1.0333 | 1.0399 | 0.93806 | 0.99721 | Dy_ppm_m163 | 0.96106 | 0.98571 | 0.91835 | 0.91178 | 0.94423 | [12 empty] | La_ppm_m139 | 0.99386 | 1.0299 | 0.99959 | 0.98976 | 1.0033 | La_ppm_m139 | 1.0143 | 1.0127 | 0.98034 | 1.0156 | 1.0057 | La_ppm_m139 | 0.96642 | 0.94963 | 1.0811 | 1.0606 | 1.0144 | Dy_ppm_m163 | 0.98078 | 0.94792 | 0.92656 | 0.93478 | 0.94751
Ce_ppm_m140 | 0.98482 | 1.0162 | 1.0437 | 0.99266 | 1.1025 | 0.96912 | 0.97501 | 0.98678 | 1.0088 | Ce_ppm_m140 | 0.97893 | 1.0005 | 1.0456 | 0.92204 | 0.98678 | Ce_ppm_m140 | 0.93773 | 0.9495 | 1.0397 | 0.92989 | 0.96422 | Ce_ppm_m140 | 0.91812 | 1.0123 | 0.95539 | 1.079 | 0.99119 | Ce_ppm_m140 | 0.95343 | 1.0927 | 1.0496 | 1.0417 | 1.0344 | Ho_ppm_m165 | 1.0096 | 0.96955 | 0.91346 | 1.0016 | 0.97356 | [7 empty] | Ce_ppm_m140 | 1.0339 | 0.94754 | 0.98482 | 0.98874 | 0.98874 | Ce_ppm_m140 | 0.99266 | 0.97501 | 1.0476 | 1.0025 | 1.0044 | Ce_ppm_m140 | 1.0103 | 1.0142 | 0.97108 | 1.0025 | 0.99953 | Ce_ppm_m140 | 0.94754 | 0.98678 | 1.0123 | 0.98089 | 1.0417 | 1.0162 | 0.99757 | Ho_ppm_m165 | 0.96154 | 1.0817 | 0.88141 | 0.94551 | 0.96755 | Ho_ppm_m165 | 0.94551 | 1.0176 | 0.96154 | 0.97756 | 0.97556 | [12 empty] | Ce_ppm_m140 | 1.0123 | 1.0064 | 1.0358 | 0.95931 | 1.0035 | Ce_ppm_m140 | 1.0299 | 1.0319 | 1.0064 | 0.9907 | 1.0147 | Ce_ppm_m140 | 1.0182 | 0.95931 | 1.0849 | 1.0299 | 1.0231 | Ho_ppm_m165 | 1.0096 | 0.97756 | 0.88141 | 0.9375 | 0.95152
Pr_ppm_m141 | 0.95776 | 1.049 | 1.0258 | 0.9918 | 1.0274 | 0.9887 | 0.95312 | 1.0042 | 0.99973 | Pr_ppm_m141 | 1.0042 | 0.99954 | 1.0289 | 0.95312 | 0.99644 | Pr_ppm_m141 | 0.93919 | 0.97323 | 0.92372 | 0.90825 | 0.9361 | Pr_ppm_m141 | 0.9361 | 0.98252 | 0.91908 | 1.0367 | 0.96859 | Pr_ppm_m141 | 0.94383 | 1.0893 | 1.0057 | 1.032 | 1.0177 | Er_ppm_m166 | 0.98468 | 0.91373 | 0.98184 | 1.0499 | 0.98255 | [7 empty] | Pr_ppm_m141 | 1.0289 | 0.98561 | 0.95002 | 0.99489 | 0.98987 | Pr_ppm_m141 | 0.9887 | 0.98561 | 1.0073 | 1.0088 | 0.9976 | Pr_ppm_m141 | 0.99489 | 0.9918 | 0.97323 | 0.96395 | 0.98097 | Pr_ppm_m141 | 0.94229 | 0.97942 | 0.99489 | 0.94693 | 1.0475 | 1.0166 | 0.98793 | Er_ppm_m166 | 1.0045 | 0.95346 | 1.0272 | 0.88252 | 0.96694 | Er_ppm_m166 | 0.98751 | 1.0102 | 1.0017 | 0.89671 | 0.97404 | [12 empty] | Pr_ppm_m141 | 1.0258 | 0.99489 | 1.046 | 0.9655 | 1.008 | Pr_ppm_m141 | 1.0011 | 1.0537 | 1.0227 | 0.97787 | 1.0138 | Pr_ppm_m141 | 0.94229 | 1.0119 | 0.94848 | 0.93765 | 0.96008 | Er_ppm_m166 | 0.92225 | 0.87401 | 0.91657 | 0.96765 | 0.92012
Nd_ppm_m146 | 0.93778 | 1.0099 | 0.9991 | 1.0135 | 1.0099 | 1.028 | 0.99188 | 0.95942 | 0.99369 | Nd_ppm_m146 | 0.98106 | 0.99549 | 0.97385 | 0.95942 | 0.97746 | Nd_ppm_m146 | 0.96303 | 0.92696 | 0.91975 | 0.91614 | 0.93147 | Nd_ppm_m146 | 0.97385 | 0.96664 | 0.96664 | 0.99188 | 0.97475 | Nd_ppm_m146 | 0.98106 | 0.99188 | 0.97024 | 0.9486 | 0.97295 | Tm_ppm_m169 | 1.0635 | 0.83811 | 0.97746 | 1.0205 | 0.9749 | [7 empty] | Nd_ppm_m146 | 1.0063 | 1.0388 | 0.99188 | 1.0712 | 1.0271 | Nd_ppm_m146 | 1.0388 | 0.97746 | 1.0027 | 0.9991 | 1.0045 | Nd_ppm_m146 | 0.96664 | 0.9486 | 0.98828 | 0.97024 | 0.96844 | Nd_ppm_m146 | 0.9486 | 0.95113 | 0.98467 | 0.97385 | 1.0243 | 1.0063 | 0.98148 | Tm_ppm_m169 | 1.2336 | 1.0635 | 1.041 | 0.90574 | 1.061 | Tm_ppm_m169 | 0.95902 | 0.9959 | 0.93033 | 1.0205 | 0.97643 | [12 empty] | Nd_ppm_m146 | 1.0063 | 0.9991 | 1.028 | 1.0388 | 1.018 | Nd_ppm_m146 | 1.0207 | 0.97746 | 1.0352 | 1.0316 | 1.0162 | Nd_ppm_m146 | 0.9486 | 1.0063 | 1.0099 | 0.93778 | 0.97565 | Tm_ppm_m169 | 0.99795 | 0.95082 | 0.93238 | 0.96311 | 0.96107
Sm_ppm_m147 | 0.99099 | 1.005 | 1.1153 | 1.0904 | 1.1743 | 0.93818 | 1.0951 | 1.0935 | 1.0628 | Sm_ppm_m147 | 0.91954 | 0.99254 | 0.98012 | 0.94905 | 0.96031 | Sm_ppm_m147 | 0.88847 | 0.87605 | 1.0221 | 0.95061 | 0.9343 | Sm_ppm_m147 | 0.95682 | 1.0593 | 1.0158 | 0.93663 | 0.99216 | Sm_ppm_m147 | 0.99565 | 1.1323 | 1.0096 | 0.97546 | 1.0283 | Yb_ppm_m172 | 0.97656 | 0.94651 | 0.97957 | 0.9375 | 0.96004 | [7 empty] | Sm_ppm_m147 | 0.92575 | 0.95216 | 1.0733 | 0.86828 | 0.95488 | Sm_ppm_m147 | 1.0252 | 1.0112 | 1.092 | 1.0158 | 1.036 | Sm_ppm_m147 | 0.93973 | 0.98322 | 0.97701 | 0.9941 | 0.97352 | Sm_ppm_m147 | 0.92886 | 0.9708 | 0.95682 | 0.91643 | 0.98478 | 0.9708 | 0.95475 | Yb_ppm_m172 | 1.1448 | 1.0637 | 0.98257 | 0.9345 | 1.0314 | Yb_ppm_m172 | 1.0577 | 0.92548 | 0.9345 | 1.0487 | 0.99159 | [12 empty] | Sm_ppm_m147 | 1.0019 | 1.0345 | 0.94284 | 1.0687 | 1.012 | Sm_ppm_m147 | 0.96925 | 0.9972 | 0.93663 | 0.93041 | 0.95837 | Sm_ppm_m147 | 0.92575 | 0.98322 | 1.0562 | 0.96925 | 0.98361 | Yb_ppm_m172 | 1.0276 | 0.9375 | 0.9345 | 1.0367 | 0.98407
Eu_ppm_m153 | 0.96404 | 0.96404 | 0.9432 | 0.92757 | 1.11 | 0.93278 | 1.0318 | 0.98489 | 0.98228 | Eu_ppm_m153 | 0.96925 | 0.94841 | 0.99531 | 0.95883 | 0.96795 | Eu_ppm_m153 | 0.91193 | 0.96404 | 1.1047 | 0.90151 | 0.97056 | Eu_ppm_m153 | 0.9432 | 1.0057 | 0.9432 | 1.0057 | 0.97447 | Eu_ppm_m153 | 1.0057 | 1.0005 | 1.1047 | 0.99531 | 1.0266 | Lu_ppm_m175 | 0.87629 | 0.85567 | 1.0722 | 0.81856 | 0.90567 | [7 empty] | Eu_ppm_m153 | 0.98489 | 0.99531 | 1.0005 | 1.0474 | 1.007 | Eu_ppm_m153 | 1.0474 | 1.0266 | 1.0318 | 1.0109 | 1.0292 | Eu_ppm_m153 | 0.93278 | 1.0057 | 0.89109 | 1.0214 | 0.96274 | Eu_ppm_m153 | 0.99531 | 0.95883 | 0.98489 | 0.92757 | 1.0005 | 1.0057 | 0.97881 | Lu_ppm_m175 | 1.1196 | 1.033 | 1.0062 | 0.90928 | 1.017 | Lu_ppm_m175 | 0.79175 | 0.90515 | 0.89485 | 0.92577 | 0.87938 | [12 empty] | Eu_ppm_m153 | 1.0214 | 0.9901 | 1.0422 | 1.0422 | 1.024 | Eu_ppm_m153 | 1.0943 | 1.0109 | 1.0214 | 0.97968 | 1.0266 | Eu_ppm_m153 | 0.9901 | 0.87024 | 1.0318 | 1.0057 | 0.97447 | Lu_ppm_m175 | 1.0784 | 0.99175 | 0.80619 | 0.89072 | 0.94175
Gd_ppm_m157 | 1.0519 | 1.0833 | 0.95126 | 0.9261 | 1.0786 | 1.0487 | 0.99686 | 1.0016 | 1.0173 | Gd_ppm_m157 | 1.011 | 1.0157 | 0.97484 | 1.022 | 1.0059 | Gd_ppm_m157 | 0.91667 | 0.98113 | 0.91509 | 0.92138 | 0.93357 | Gd_ppm_m157 | 1.0126 | 0.98428 | 0.91509 | 1.0299 | 0.98546 | Gd_ppm_m157 | 0.95755 | 0.91824 | 0.94497 | 0.99371 | 0.95362 | Hf_ppm_m177 | 1.0061 | 0.96166 | 0.98072 | 1.0104 | 0.98973 | [7 empty] | Gd_ppm_m157 | 0.86006 | 0.81604 | 0.9434 | 1.066 | 0.92138 | Gd_ppm_m157 | 1.0283 | 1.0142 | 1.0991 | 1.0456 | 1.0468 | Gd_ppm_m157 | 1.0267 | 0.95126 | 1.0236 | 0.98428 | 0.99646 | Gd_ppm_m157 | 1.0031 | 0.96855 | 0.97327 | 0.93553 | 1.0519 | 0.93868 | 0.97851 | Hf_ppm_m177 | 1.1714 | 1.057 | 0.98708 | 1.0019 | 1.0543 | Hf_ppm_m177 | 1.0146 | 1.0167 | 0.97013 | 0.96802 | 0.99237 | [12 empty] | Gd_ppm_m157 | 1.0739 | 0.97799 | 1.055 | 0.99843 | 1.0263 | Gd_ppm_m157 | 1.0236 | 0.99371 | 1.0314 | 1.0267 | 1.0189 | Gd_ppm_m157 | 1.0314 | 1.0377 | 0.94654 | 0.97799 | 0.99843 | Hf_ppm_m177 | 0.91718 | 0.97649 | 0.92353 | 0.94048 | 0.93942
Tb_ppm_m159 | 0.99288 | 1.0315 | 1.0285 | 1.0437 | 0.92879 | 1.0458 | 1.0519 | 1.0682 | 1.0239 | Tb_ppm_m159 | 0.97558 | 0.98881 | 0.97762 | 0.93082 | 0.96821 | Tb_ppm_m159 | 0.96643 | 0.90437 | 1.0417 | 0.9766 | 0.97228 | Tb_ppm_m159 | 1.0173 | 0.94608 | 0.88505 | 1.0071 | 0.96389 | Tb_ppm_m159 | 1.0275 | 1.058 | 1.0102 | 1.0285 | 1.031 | Ta_ppm_m181 | 1.0278 | 1.0833 | 1.0278 | 1.0417 | 1.0451 | [7 empty] | Tb_ppm_m159 | 0.95015 | 0.98067 | 1.0478 | 1.0376 | 1.0041 | Tb_ppm_m159 | 1.06 | 1.0783 | 0.96033 | 1.0275 | 1.0315 | Tb_ppm_m159 | 0.9766 | 0.89725 | 1.061 | 0.93388 | 0.96719 | Tb_ppm_m159 | 0.941 | 1.0031 | 0.95524 | 0.95829 | 1.0621 | 0.97355 | 0.9822 | Ta_ppm_m181 | 1.0194 | 1.0694 | 0.975 | 0.93611 | 1 | Ta_ppm_m181 | 0.95833 | 0.95833 | 0.94444 | 0.97639 | 0.95937 | [12 empty] | Tb_ppm_m159 | 1.0387 | 1.0366 | 1.0773 | 1.0458 | 1.0496 | Tb_ppm_m159 | 0.99695 | 1.0346 | 0.99797 | 1.1089 | 1.0346 | Tb_ppm_m159 | 0.94608 | 1.0173 | 1.0682 | 0.98678 | 1.0046 | Ta_ppm_m181 | 0.95833 | 0.90694 | 1.0972 | 1.0306 | 0.99826
Dy_ppm_m163 | 0.96106 | 1.0416 | 1.1089 | 1.0826 | 1.0547 | 1.0038 | 1.0366 | 1.035 | 1.0405 | Dy_ppm_m163 | 0.94135 | 1.0268 | 0.97421 | 0.97749 | 0.97996 | Dy_ppm_m163 | 0.97749 | 0.99885 | 0.95778 | 0.96435 | 0.97462 | Dy_ppm_m163 | 0.96271 | 1.1336 | 1.0284 | 0.92656 | 1.0128 | Dy_ppm_m163 | 0.87892 | 1.0793 | 0.99228 | 1.0186 | 0.99228 | W_ppm_m182 | 1.19 | 1.0557 | 1.094 | 0.97889 | 1.0797 | [7 empty] | Dy_ppm_m163 | 1.0662 | 0.94792 | 1.0646 | 0.99228 | 1.0177 | Dy_ppm_m163 | 1.0366 | 0.99885 | 0.94628 | 0.97256 | 0.98858 | Dy_ppm_m163 | 0.97421 | 1.0186 | 0.96928 | 1.0399 | 1.0005 | Dy_ppm_m163 | 1.0021 | 0.96928 | 1.0153 | 0.93478 | 0.96928 | 1.0021 | 0.98215 | W_ppm_m182 | 1.4395 | 0.92131 | 1.0173 | 0.9405 | 1.0797 | W_ppm_m182 | 0.92131 | 0.97889 | 1.094 | 1.0749 | 1.0173 | [12 empty] | Dy_ppm_m163 | 1.0268 | 1.058 | 1.012 | 1.0103 | 1.0268 | Dy_ppm_m163 | 1.0038 | 0.97914 | 1.012 | 0.97092 | 0.99146 | Dy_ppm_m163 | 0.99064 | 0.94792 | 1.1056 | 1.0366 | 1.0202 | W_ppm_m182 | 1.19 | 1.0365 | 1.144 | 1.0173 | 1.0969
Ho_ppm_m165 | 0.95833 | 1.0505 | 1.0096 | 0.89103 | 1.0337 | 0.96955 | 0.97756 | 0.96154 | 0.98147 | Ho_ppm_m165 | 1.0008 | 0.98317 | 1.0409 | 0.91106 | 0.98397 | Ho_ppm_m165 | 0.91907 | 0.99359 | 1.0096 | 0.89744 | 0.95493 | Ho_ppm_m165 | 0.96154 | 0.95353 | 0.90545 | 1.0176 | 0.95954 | Ho_ppm_m165 | 0.92949 | 1.1538 | 1.0873 | 0.98558 | 1.0391 | Tl_ppm_m205 | 1.2388 | 0.8209 | 1.0547 | 1.2587 | 1.0933 | [7 empty] | Ho_ppm_m165 | 1.0417 | 1.0016 | 0.99359 | 1.0897 | 1.0317 | Ho_ppm_m165 | 0.97756 | 0.88141 | 1.1058 | 0.98558 | 0.98758 | Ho_ppm_m165 | 1.0337 | 0.94391 | 0.92228 | 0.97596 | 0.96895 | Ho_ppm_m165 | 0.88622 | 1.0312 | 1.0128 | 0.92228 | 1.0737 | 1.0321 | 0.99306 | Tl_ppm_m205 | 0.8408 | 0.99502 | 1.3085 | 1.3284 | 1.1182 | Tl_ppm_m205 | 1.1244 | 1.1692 | 0.86567 | 1.1592 | 1.0796 | [12 empty] | Ho_ppm_m165 | 1.0312 | 1.0192 | 1.0553 | 1.0433 | 1.0373 | Ho_ppm_m165 | 1.0986 | 1.0048 | 0.94471 | 1.0256 | 1.0184 | Ho_ppm_m165 | 0.95353 | 1.0577 | 0.95353 | 0.97756 | 0.98558 | Tl_ppm_m205 | 0.64179 | 0.86567 | 0.78607 | 0.77114 | 0.76617
Er_ppm_m166 | 0.979 | 1.0358 | 0.96481 | 1.0301 | 1.0131 | 1.0244 | 1.0159 | 1.0414 | 1.0131 | Er_ppm_m166 | 1.0187 | 1.0159 | 1.067 | 0.97049 | 1.018 | Er_ppm_m166 | 0.94211 | 1.0443 | 0.95914 | 0.87117 | 0.95417 | Er_ppm_m166 | 1.0017 | 0.93644 | 0.84279 | 1.0244 | 0.95133 | Er_ppm_m166 | 0.98468 | 0.99886 | 0.979 | 1.0074 | 0.99248 | Pb_ppm_m208 | 1.1728 | 1.1236 | 1.1176 | 1.1658 | 1.1449 | [7 empty] | Er_ppm_m166 | 1.0131 | 1.1209 | 1.0216 | 1.0868 | 1.0606 | Er_ppm_m166 | 1.0017 | 0.96481 | 1.1039 | 0.979 | 1.0123 | Er_ppm_m166 | 1.0556 | 0.95346 | 1.0556 | 1.084 | 1.0372 | Er_ppm_m166 | 0.92225 | 0.94211 | 0.9336 | 0.98468 | 1.0528 | 0.97049 | 0.96765 | Pb_ppm_m208 | 0.95692 | 0.95291 | 1.0302 | 0.97098 | 0.97776 | Pb_ppm_m208 | 1.0011 | 1.0212 | 0.92078 | 1.0302 | 0.99332 | [12 empty] | Er_ppm_m166 | 1.0272 | 1.0159 | 1.0131 | 0.97616 | 1.0081 | Er_ppm_m166 | 1.0925 | 0.98751 | 1.0556 | 1.0499 | 1.0464 | Er_ppm_m166 | 0.87684 | 0.99603 | 1.067 | 0.96765 | 0.97687 | Pb_ppm_m208 | 0.97701 | 0.93985 | 0.91073 | 0.9509 | 0.94462
Tm_ppm_m169 | 1.0902 | 1.0348 | 0.9959 | 0.98566 | 1.0082 | 1.041 | 0.98156 | 1.0451 | 1.0228 | Tm_ppm_m169 | 0.94262 | 1.0615 | 1.0984 | 1.0369 | 1.0348 | Tm_ppm_m169 | 0.90164 | 0.91598 | 1.0984 | 0.93852 | 0.96363 | Tm_ppm_m169 | 1.0348 | 1.0266 | 1.0205 | 0.85861 | 0.98514 | Tm_ppm_m169 | 0.95697 | 1.0963 | 0.98156 | 0.95492 | 0.99744 | Th_ppm_m232 | 1.0017 | 1.0362 | 0.98103 | 1.0103 | 1.0073 | [7 empty] | Tm_ppm_m169 | 1.0041 | 0.94467 | 0.8873 | 1.0615 | 0.97439 | Tm_ppm_m169 | 0.92213 | 0.89959 | 0.91803 | 1.0164 | 0.93904 | Tm_ppm_m169 | 1.0328 | 0.96107 | 0.93852 | 0.9959 | 0.98207 | Tm_ppm_m169 | 1.0205 | 0.97746 | 0.92623 | 0.98566 | 1.0328 | 0.93033 | 0.97883 | Th_ppm_m232 | 1.0069 | 0.97414 | 0.95345 | 0.95 | 0.97112 | Th_ppm_m232 | 0.92586 | 0.96379 | 0.95172 | 0.97586 | 0.95431 | [12 empty] | Tm_ppm_m169 | 1.0615 | 1.0717 | 0.9918 | 0.99385 | 1.0297 | Tm_ppm_m169 | 1.1148 | 1.041 | 1.0123 | 1.0184 | 1.0466 | Tm_ppm_m169 | 1.0246 | 1.002 | 1.2295 | 1.1045 | 1.0902 | Th_ppm_m232 | 1.0172 | 0.9569 | 0.94138 | 0.99483 | 0.97759
Yb_ppm_m172 | 1.0126 | 0.95553 | 0.96755 | 1.0276 | 1.1599 | 0.98858 | 1.1298 | 1.0877 | 1.0412 | Yb_ppm_m172 | 0.98558 | 0.96755 | 0.94952 | 0.99459 | 0.97431 | Yb_ppm_m172 | 0.93149 | 1.0427 | 1.0006 | 0.83534 | 0.95252 | Yb_ppm_m172 | 0.88942 | 1.0487 | 0.9405 | 1.0246 | 0.97581 | Yb_ppm_m172 | 0.79627 | 1.1058 | 1.0337 | 0.80829 | 0.936 | U_ppm_m238 | 0.98256 | 1.0291 | 0.93605 | 0.91279 | 0.96512 | [7 empty] | Yb_ppm_m172 | 0.97656 | 0.90144 | 1.1058 | 1.0577 | 1.0104 | Yb_ppm_m172 | 0.96454 | 0.8774 | 1.1058 | 1.0667 | 1.0036 | Yb_ppm_m172 | 0.91346 | 0.9976 | 0.91346 | 0.98858 | 0.95328 | Yb_ppm_m172 | 0.95252 | 1.0517 | 1.0066 | 0.95553 | 0.9345 | 0.97055 | 0.97857 | U_ppm_m238 | 1.0698 | 0.97674 | 0.99419 | 1.0058 | 1.0116 | U_ppm_m238 | 0.94767 | 0.98837 | 1.0233 | 0.98837 | 0.98692 | [12 empty] | Yb_ppm_m172 | 0.99159 | 0.95553 | 1.0306 | 1.0156 | 0.99835 | Yb_ppm_m172 | 1.0216 | 0.99159 | 1.0036 | 1.0637 | 1.0201 | Yb_ppm_m172 | 0.97055 | 1.0487 | 1.1118 | 1.0306 | 1.0404 | U_ppm_m238 | 1.1628 | 1.0407 | 1.1512 | 1.0291 | 1.0959
Lu_ppm_m175 | 0.96289 | 0.89691 | 1.0165 | 0.97732 | 1.0082 | 1.0021 | 1.0701 | 0.99588 | 0.99124 | Lu_ppm_m175 | 0.89485 | 1.0289 | 1.0887 | 0.98969 | 1.0005 | Lu_ppm_m175 | 0.84948 | 1.1546 | 1.0557 | 0.72371 | 0.94588 | Lu_ppm_m175 | 1.0309 | 0.98763 | 0.93814 | 1.0103 | 0.99175 | Lu_ppm_m175 | 1.0412 | 1.1485 | 0.98763 | 0.91546 | 1.0232 | [13 empty] | Lu_ppm_m175 | 1.0701 | 1.0454 | 0.96289 | 1.0495 | 1.032 | Lu_ppm_m175 | 1.0124 | 1.0309 | 0.98557 | 1.0124 | 1.0103 | Lu_ppm_m175 | 0.90722 | 0.96289 | 0.96495 | 1.0412 | 0.96907 | Lu_ppm_m175 | 0.94227 | 0.95052 | 1.068 | 0.92165 | 1.0351 | 1.0515 | 0.99485 | [24 empty] | Lu_ppm_m175 | 1.0268 | 1.0351 | 1.1031 | 0.98557 | 1.0376 | Lu_ppm_m175 | 1.134 | 1.0866 | 1.0454 | 0.99381 | 1.0649 | Lu_ppm_m175 | 0.98557 | 1.0557 | 0.96495 | 1.0454 | 1.0129
[34 empty] | Th/U | 1.0195 | 1.0069 | 1.0481 | 1.1069 | 1.0437 | [33 empty] | Th/U | 0.94123 | 0.99733 | 0.95902 | 0.94451 | 0.95996 | Th/U | 0.97698 | 0.97513 | 0.93009 | 0.98734 | 0.96696 | [30 empty] | Th/U | 0.87483 | 0.91948 | 0.81776 | 0.96673 | 0.89202
Hf_ppm_m177 | 0.95319 | 1.1311 | 1.0294 | 0.95531 | 0.9892 | 1.0019 | 0.90023 | 0.95742 | 0.98973 | Hf_ppm_m177 | 1.0443 | 0.95319 | 1.0019 | 0.97649 | 0.99396 | Hf_ppm_m177 | 1.0655 | 0.98072 | 0.90871 | 0.99132 | 0.98655 | Hf_ppm_m177 | 1.0379 | 1.0527 | 0.97649 | 1.0104 | 1.0194 | Hf_ppm_m177 | 0.98072 | 0.93412 | 0.99132 | 0.96378 | 0.96749 | [13 empty] | Hf_ppm_m177 | 0.85999 | 0.99979 | 0.96378 | 1.0506 | 0.96854 | Hf_ppm_m177 | 1.04 | 0.96802 | 1.0506 | 1.1015 | 1.04 | Hf_ppm_m177 | 0.9892 | 0.9659 | 0.91718 | 1.0443 | 0.97914 | Hf_ppm_m177 | 0.99555 | 0.98496 | 1.0083 | 0.99343 | 1.0527 | 1.0591 | 1.0157 | [24 empty] | Hf_ppm_m177 | 1.021 | 1.0019 | 1.0782 | 0.96378 | 1.0162 | Hf_ppm_m177 | 1.0845 | 1.0422 | 1.0104 | 0.97225 | 1.0273 | Hf_ppm_m177 | 1.004 | 1.0527 | 1.0316 | 0.88964 | 0.99449
Ta_ppm_m181 | 0.97639 | 1.025 | 1.0611 | 1.0625 | 0.93333 | 1.0833 | 1.0569 | 0.97083 | 1.0212 | Ta_ppm_m181 | 0.9 | 1.0611 | 1 | 0.94583 | 0.97674 | Ta_ppm_m181 | 0.99167 | 0.94444 | 1.0694 | 0.88889 | 0.97361 | Ta_ppm_m181 | 0.90278 | 0.95833 | 0.80694 | 1.0694 | 0.93437 | Ta_ppm_m181 | 0.97222 | 0.96528 | 1.0444 | 0.96111 | 0.98576 | [13 empty] | Ta_ppm_m181 | 0.78472 | 1.0556 | 0.9375 | 0.96667 | 0.93611 | Ta_ppm_m181 | 1.0472 | 1.0639 | 0.92639 | 0.82083 | 0.96458 | Ta_ppm_m181 | 0.92778 | 0.93056 | 0.93889 | 1.0681 | 0.96632 | Ta_ppm_m181 | 0.9125 | 0.92222 | 0.94306 | 0.8875 | 0.95833 | 1.0097 | 0.93889 | [24 empty] | Ta_ppm_m181 | 0.99444 | 1.0778 | 1.0292 | 1.0028 | 1.026 | Ta_ppm_m181 | 0.97778 | 0.91667 | 0.94306 | 1.0458 | 0.97083 | Ta_ppm_m181 | 1 | 0.85556 | 1.1111 | 1.1486 | 1.0288
W_ppm_m182 | 0.83301 | 0.91555 | 1.1593 | 0.91555 | 1.2745 | 0.96161 | 1.1132 | 1.0422 | 1.0269 | W_ppm_m182 | 0.91171 | 1.0422 | 1.1344 | 0.93474 | 1.0058 | W_ppm_m182 | 0.92131 | 1.0173 | 0.95969 | 1.0557 | 0.98848 | W_ppm_m182 | 0.90211 | 0.72937 | 0.99808 | 0.74856 | 0.84453 | W_ppm_m182 | 0.87908 | 1.0173 | 0.98848 | 0.91555 | 0.9501 | [13 empty] | W_ppm_m182 | 1.0173 | 1.0173 | 1.1324 | 1.094 | 1.0653 | W_ppm_m182 | 1.1324 | 1.0557 | 1.286 | 1.19 | 1.166 | W_ppm_m182 | 0.98273 | 0.90787 | 1.1305 | 1.048 | 1.0173 | W_ppm_m182 | 0.99616 | 1.0345 | 1.0883 | 0.87908 | 1.094 | 1.0576 | 1.025 | [24 empty] | W_ppm_m182 | 0.92706 | 0.94818 | 1.0307 | 0.89827 | 0.95106 | W_ppm_m182 | 1.0077 | 0.8906 | 1.1324 | 1.0307 | 1.0154 | W_ppm_m182 | 1.0365 | 1.0557 | 1.0557 | 0.96737 | 1.0288
Tl_ppm_m205 | 1.2985 | 1.1443 | 1.0299 | 0.9204 | 0.72637 | 1.0945 | 1.0249 | 0.91045 | 1.0187 | Tl_ppm_m205 | 1.3383 | 1.2786 | 1.2338 | 1.1791 | 1.2575 | Tl_ppm_m205 | 1.2438 | 1.3035 | 1.1841 | 1.2189 | 1.2376 | Tl_ppm_m205 | 0.91045 | 1.4428 | 1.3731 | 1.3532 | 1.2699 | Tl_ppm_m205 | 1.1095 | 1.0149 | 1.2935 | 1.3284 | 1.1866 | [13 empty] | Tl_ppm_m205 | 0.9801 | 0.93035 | 0.86567 | 1.2537 | 1.0075 | Tl_ppm_m205 | 1.2985 | 1.5522 | 1.398 | 0.9403 | 1.2973 | Tl_ppm_m205 | 1.0746 | 0.96517 | 0.99005 | 0.8408 | 0.96766 | Tl_ppm_m205 | 0.76617 | 0.71144 | 0.71642 | 0.9204 | 1.2189 | 1.199 | 0.92206 | [24 empty] | Tl_ppm_m205 | 0.89055 | 1.0498 | 0.90547 | 1.0398 | 0.97139 | Tl_ppm_m205 | 1.2736 | 1.0547 | 1.0249 | 1.0498 | 1.1007 | Tl_ppm_m205 | 0.75622 | 0.69652 | 0.78109 | 0.81095 | 0.76119
Pb_ppm_m208 | 0.99709 | 0.98705 | 1.0292 | 1.0393 | 0.98604 | 1.0091 | 1.0413 | 1.0031 | 1.0115 | Pb_ppm_m208 | 1.0373 | 1.0784 | 1.1146 | 1.0021 | 1.0581 | Pb_ppm_m208 | 0.97701 | 1.0573 | 1.1377 | 1.0071 | 1.0448 | Pb_ppm_m208 | 1.0292 | 1.0242 | 1.1156 | 1.1818 | 1.0877 | Pb_ppm_m208 | 1.0513 | 1.2321 | 1.1266 | 1.0875 | 1.1244 | [13 empty] | Pb_ppm_m208 | 1.0162 | 0.93383 | 0.95893 | 0.99006 | 0.97475 | Pb_ppm_m208 | 1.0724 | 1.0242 | 1.0011 | 0.9278 | 1.0064 | Pb_ppm_m208 | 0.99307 | 0.97098 | 0.90471 | 0.95592 | 0.95617 | Pb_ppm_m208 | 0.96295 | 0.93383 | 0.94889 | 0.96696 | 1.0413 | 1.0473 | 0.98353 | [24 empty] | Pb_ppm_m208 | 0.95692 | 1.0573 | 1.0634 | 0.99408 | 1.0179 | Pb_ppm_m208 | 1.0362 | 1.0152 | 0.99608 | 0.96797 | 1.0039 | Pb_ppm_m208 | 0.94588 | 0.95391 | 0.94789 | 0.92881 | 0.94412
Bi_ppm_m209 | 0.94885 | 0.80551 | 0.66621 | 0.90847 | 1.6756 | 0.94885 | 0.62584 | 1.2315 | 0.97636 | Bi_ppm_m209 | 0.76715 | 0.88828 | 0.8681 | 0.74697 | 0.81763 | Bi_ppm_m209 | 0.88828 | 1.1709 | 0.72678 | 0.88828 | 0.91857 | Bi_ppm_m209 | 1.2719 | 0.66621 | 1.1104 | 1.5545 | 1.1507 | Bi_ppm_m209 | 1.3526 | 1.1507 | 0.76715 | 0.88828 | 1.0397 | [13 empty] | Bi_ppm_m209 | 1.3728 | 1.6353 | 1.2719 | 1.1507 | 1.3577 | Bi_ppm_m209 | 0.70659 | 1.1104 | 2.0592 | 1.4536 | 1.3324 | Bi_ppm_m209 | 1.1507 | 1.393 | 1.6151 | 1.1507 | 1.3274 | Bi_ppm_m209 | 1.3728 | 1.2517 | 1.4536 | 0.82772 | 0.84791 | 0.92866 | 1.1137 | [24 empty] | Bi_ppm_m209 | 1.3122 | 1.0498 | 0.80753 | 0.60363 | 0.9433 | Bi_ppm_m209 | 0.74697 | 0.96904 | 0.64602 | 0.88828 | 0.81258 | Bi_ppm_m209 | 1.4132 | 1.1507 | 1.292 | 1.6958 | 1.3879
Th_ppm_m232 | 1.0069 | 0.99828 | 1.0414 | 1.0759 | 1.0345 | 1.0379 | 1.0034 | 1.0534 | 1.0315 | Th_ppm_m232 | 1.0086 | 1.0086 | 0.99483 | 0.96552 | 0.9944 | Th_ppm_m232 | 0.92069 | 0.93621 | 1.0517 | 0.88966 | 0.94957 | Th_ppm_m232 | 0.93621 | 1.0621 | 1.019 | 1.0207 | 1.0095 | Th_ppm_m232 | 0.91207 | 1.019 | 0.94655 | 0.97414 | 0.96293 | [13 empty] | Th_ppm_m232 | 0.98103 | 0.9931 | 0.93621 | 0.98793 | 0.97457 | Th_ppm_m232 | 1.069 | 0.97759 | 1.0466 | 1.0293 | 1.0306 | Th_ppm_m232 | 0.99138 | 1.0138 | 0.93966 | 1.0103 | 0.98879 | Th_ppm_m232 | 0.97241 | 0.95517 | 0.98448 | 0.96207 | 0.98448 | 1.019 | 0.9796 | [24 empty] | Th_ppm_m232 | 1.0155 | 1.0724 | 1.069 | 0.99483 | 1.0379 | Th_ppm_m232 | 1.0293 | 1.0103 | 0.99655 | 1.0431 | 1.0198 | Th_ppm_m232 | 0.97414 | 0.93966 | 1.0845 | 1.0293 | 1.0069
U_ppm_m238 | 0.94186 | 0.97093 | 0.93605 | 0.94767 | 1.186 | 1.0116 | 0.90116 | 1.0058 | 0.98765 | U_ppm_m238 | 0.9343 | 0.95988 | 0.99942 | 0.90116 | 0.94869 | U_ppm_m238 | 0.87209 | 0.86047 | 1.1628 | 0.9186 | 0.95349 | U_ppm_m238 | 0.8314 | 0.97093 | 0.94186 | 1.0698 | 0.95349 | U_ppm_m238 | 0.90698 | 1.0756 | 1.0174 | 0.94186 | 0.98547 | [13 empty] | U_ppm_m238 | 0.99419 | 1.0174 | 1.0465 | 1.0058 | 1.016 | U_ppm_m238 | 1.0116 | 1.0058 | 1.2035 | 1.1047 | 1.0814 | U_ppm_m238 | 1.0698 | 1.0814 | 1.064 | 1.093 | 1.077 | U_ppm_m238 | 0.98256 | 0.99244 | 1.093 | 0.97093 | 0.98256 | 1 | 1.0036 | [24 empty] | U_ppm_m238 | 1.011 | 1.0116 | 0.98663 | 0.96279 | 0.99302 | U_ppm_m238 | 1.0523 | 1 | 0.94186 | 1.0233 | 1.0044 | U_ppm_m238 | 1.0116 | 1.0349 | 1.0698 | 1.0756 | 1.048
[10 empty] | Th/U | 1.0795 | 1.0508 | 0.99541 | 1.0714 | 1.0482 | Th/U | 1.0557 | 1.088 | 0.90448 | 0.96849 | 0.99589 | Th/U | 1.1261 | 1.0939 | 1.0819 | 0.95412 | 1.0587 | Th/U | 1.0056 | 0.94736 | 0.93033 | 1.0343 | 0.97713 | [13 empty] | Th/U | 0.98677 | 0.97608 | 0.8946 | 0.98222 | 0.95923 | Th/U | 1.0567 | 0.97194 | 0.8696 | 0.9318 | 0.95303 | Th/U | 0.92672 | 0.93749 | 0.88317 | 0.92436 | 0.91807 | Th/U | 0.98968 | 0.96245 | 0.9007 | 0.99087 | 1.002 | 1.019 | 0.9761 | [30 empty] | Th/U | 0.97813 | 1.0103 | 1.0581 | 1.0194 | 1.0154
[103 empty] | Th/Ta | 1.0527 | 1.1022 | 1.0567 | 0.99739 | 1.0505

## Preferred_Values
Numbers on "Reduced to Ca" and "Si" are shown as Measured/Preferred
Pref Values
BCR-2G
Li_ppm_m7 | 9.374
Be_ppm_m9 | 2.097
Na_ppm_m23 | 23962
Mg_ppm_m24 | 21470
Al_ppm_m27 | 70920
P_ppm_m31 | 1615
K_ppm_m39 | 14900
Ca_ppm_m43 | 7.3673
Sc_ppm_m45 | 34.046
Ti_ppm_m49 | 14057.1924
V_ppm_m51 | 436.642
Cr_ppm_m53 | 16.065
Mn_ppm_m55 | 1630
Fe_ppm_m56 | 93600
Co_ppm_m59 | 38.11
Ni_ppm_m60 | 12.224
Cu_ppm_m65 | 18.358
Zn_ppm_m66 | 154.298
Ga_ppm_m71 | 21.537
Ge_ppm_m74 | 1.81
Rb_ppm_m85 | 48.09
Sr_ppm_m88 | 340.328
Y_ppm_m89 | 32.605
Zr_ppm_m90 | 175.087
Nb_ppm_m93 | 12.117
Mo_ppm_m95 | 250.491
Cd_ppm_m111 | 0.20343
In_ppm_m115 | 0.098043
Sn_ppm_m118 | 2.074
Sb_ppm_m121 | 0.321
Cs_ppm_m133 | 1.179
Ba_ppm_m137 | 670
La_ppm_m139 | 24.42
Ce_ppm_m140 | 50.974
Pr_ppm_m141 | 6.463
Nd_ppm_m146 | 27.725
Sm_ppm_m147 | 6.438
Eu_ppm_m153 | 1.919
Gd_ppm_m157 | 6.36
Tb_ppm_m159 | 0.983
Dy_ppm_m163 | 6.087
Ho_ppm_m165 | 1.248
Er_ppm_m166 | 3.524
Tm_ppm_m169 | 0.488
Yb_ppm_m172 | 3.328
Lu_ppm_m175 | 0.485
Hf_ppm_m177 | 4.721
Ta_ppm_m181 | 0.72
W_ppm_m182 | 0.521
Tl_ppm_m205 | 0.201
Pb_ppm_m208 | 9.959
Bi_ppm_m209 | 0.049534
Th_ppm_m232 | 5.8
U_ppm_m238 | 1.72

## Reduced to Si
Glasses 29/10/2019 Batch 2 30 um | [5 empty] | MIs 30/10/2019 Batch 3 50 um
Reduced to Si | [5 empty] | Reduced to Si
| BCR_30um_1 | BCR_30um_2 | BCR_30um_3 | BCR_30um_4 | AVERAGE |  | BCR_50um_9 | BCR_50um_10 | BCR_50um_11 | BCR_50um_12 | BCR_50um_13 | BCR_50um_14 | AVERAGE
Li_ppm_m7 | 1.1308 | 1.0134 | 0.97077 | 1.0134 | 1.0321 | Li_ppm_m7 | 0.91636 | 1.0156 | 1.038 | 0.94837 | 0.9825 | 0.98784 | 0.98144
Be_ppm_m9 | 0.86791 | 0.64855 | 0.79161 | 0.96328 | 0.81784 | Be_ppm_m9 | 1.154 | 0.77253 | 0.99666 | 1.0968 | 0.61993 | 1.1111 | 0.95851
Na_ppm_m23 | 0.94316 | 1.1017 | 1.0183 | 1.1143 | 1.0444 | Na_ppm_m23 | 0.99324 | 1.0099 | 1.0433 | 0.97571 | 0.99324 | 1.0099 | 1.0042
K_ppm_m39 | 0.98658 | 0.91812 | 0.83893 | 0.87248 | 0.90403 | K_ppm_m39 | 0.93758 | 0.93624 | 0.98926 | 0.88591 | 0.9302 | 0.97651 | 0.94262
Ca_ppm_m43 | 0.94065 | 0.95151 | 0.89857 | 0.98001 | 9.4269e-05 | Ca_ppm_m43 | 0.91486 | 0.96101 | 0.98816 | 0.89993 | 0.89857 | 0.923 | 0.93092
Sc_ppm_m45 | 0.93403 | 0.99277 | 1.0075 | 1.0045 | 0.9847 | Sc_ppm_m45 | 0.9869 | 0.98984 | 0.97515 | 0.95459 | 0.93697 | 0.93109 | 0.96242
Ti_ppm_m49 | 0.88211 | 0.89634 | 0.96392 | 1.0173 | 0.93991 | Ti_ppm_m49 | 0.87998 | 0.87429 | 0.94044 | 0.93048 | 0.9938 | 0.97744 | 0.93274
V_ppm_m51 | 0.99395 | 0.92753 | 0.93669 | 1.0146 | 0.96818 | V_ppm_m51 | 0.97105 | 0.95502 | 0.97792 | 0.94585 | 0.97563 | 0.95273 | 0.96303
Cr_ppm_m53 | 1.0769 | 1.0955 | 0.84656 | 0.87768 | 0.97417 | Cr_ppm_m53 | 0.90881 | 1.0395 | 1.1267 | 1.0146 | 0.97728 | 0.97728 | 1.0074
Mn_ppm_m55 | 0.98773 | 0.96933 | 0.88957 | 0.95092 | 0.94939 | Mn_ppm_m55 | 0.93742 | 0.9638 | 0.97055 | 0.8865 | 0.9227 | 0.95092 | 0.93865
Fe_ppm_m56 | 0.99786 | 1.0897 | 1.0662 | 1.0577 | 1.0529 | Fe_ppm_m56 | 1.062 | 1.1186 | 1.0951 | 1.0053 | 0.98718 | 1.0524 | 1.0534
Co_ppm_m59 | 0.92889 | 1.0706 | 1.0155 | 1.0155 | 1.0076 | Co_ppm_m59 | 1.026 | 1.0365 | 1.026 | 0.98399 | 0.95513 | 0.99974 | 1.0045
Ni_ppm_m60 | 0.90805 | 1.0798 | 1.0962 | 1.1126 | 1.0492 | Ni_ppm_m60 | 1.0144 | 0.91623 | 0.95713 | 0.9375 | 0.97349 | 0.94077 | 0.95659
Cu_ppm_m65 | 0.93692 | 0.94782 | 0.95326 | 0.94237 | 0.94509 | Cu_ppm_m65 | 0.9696 | 0.93147 | 1.0132 | 0.92603 | 1.0132 | 0.91513 | 0.96143
Zn_ppm_m66 | 1.0758 | 0.9527 | 0.97214 | 1.0175 | 1.0045 | Zn_ppm_m66 | 0.92354 | 1.0052 | 0.99288 | 0.9365 | 1.011 | 0.9203 | 0.96491
Ga_ppm_m71 | 1.119 | 0.96114 | 0.91935 | 0.97042 | 0.99248 | Ga_ppm_m71 | 0.95185 | 1.0447 | 1.0638 | 1.0011 | 1.0029 | 1.0122 | 1.0128
Ge_ppm_m74 | 0.8895 | 1.0442 | 0.69613 | 0.77348 | 0.85083 | Ge_ppm_m74 | 0.8011 | 0.83425 | 0.88398 | 0.75138 | 0.81768 | 0.8011 | 0.81492
Rb_ppm_m85 | 1.0189 | 1.0834 | 1.0064 | 0.95238 | 1.0153 | Rb_ppm_m85 | 1.021 | 1.046 | 1.046 | 0.99813 | 0.96902 | 1.0002 | 1.0134
Sr_ppm_m88 | 0.94321 | 0.96378 | 0.96084 | 1.0431 | 0.97773 | Sr_ppm_m88 | 0.98434 | 0.97847 | 0.96378 | 0.9579 | 0.93145 | 0.96084 | 0.9628
Y_ppm_m89 | 0.97224 | 0.95077 | 0.91397 | 1.0857 | 0.98068 | Y_ppm_m89 | 0.98758 | 0.96611 | 0.97531 | 0.94464 | 0.95077 | 0.95691 | 0.96355
Zr_ppm_m90 | 1.0395 | 0.9481 | 0.93668 | 1.0566 | 0.99522 | Zr_ppm_m90 | 0.96409 | 0.94696 | 0.97894 | 0.9561 | 0.92754 | 0.96523 | 0.95648
Nb_ppm_m93 | 1.0068 | 0.96559 | 0.89956 | 0.94083 | 0.95321 | Nb_ppm_m93 | 0.93422 | 1.0035 | 0.98292 | 0.93918 | 0.9433 | 0.95981 | 0.9605
Mo_ppm_m95 | 1.042 | 1.1378 | 1.018 | 1.018 | 1.0539 | Mo_ppm_m95 | 1.03 | 1.1178 | 1.1018 | 1.0228 | 1.034 | 1.0619 | 1.0614
Cd_ppm_m111 | #VALUE! | #VALUE! | #VALUE! | 0.68819 | 0.68819 | Cd_ppm_m111 | 0.68819 | #VALUE! | 1.2781 | 0.86516 | 0.83566 | 1.1306 | 0.95954
In_ppm_m115 | 1.2342 | 0.87716 | 1.4279 | 1.5197 | 1.2647 | In_ppm_m115 | 1.0506 | 1.1424 | 1.2749 | 1.071 | 0.99956 | 1.1832 | 1.1203
Sn_ppm_m118 | 1.1427 | 0.95468 | 0.86789 | 0.94986 | 0.97878 | Sn_ppm_m118 | 1.08 | 0.97396 | 1.0704 | 0.91128 | 1.027 | 1.0415 | 1.0174
Sb_ppm_m121 | 0.96262 | 1.215 | 1.0125 | 0.85981 | 1.0125 | Sb_ppm_m121 | 1.9564 | 0.87227 | 0.93769 | 0.96573 | 0.94704 | 1.0031 | 1.1137
Cs_ppm_m133 | 0.9754 | 0.98388 | 0.91603 | 1.0093 | 0.97116 | Cs_ppm_m133 | 1.0136 | 1.0161 | 1.0136 | 0.97455 | 1.0093 | 1.0085 | 1.0059
Ba_ppm_m137 | 0.93731 | 1.0418 | 0.9194 | 0.96567 | 0.96604 | Ba_ppm_m137 | 0.97164 | 0.99403 | 1.0224 | 0.96418 | 0.96418 | 0.99552 | 0.98532
La_ppm_m139 | 0.92138 | 0.96642 | 1.0483 | 1.0197 | 0.98894 | La_ppm_m139 | 0.94595 | 0.92547 | 0.95823 | 0.94799 | 0.97256 | 0.97052 | 0.95345
Ce_ppm_m140 | 1.0044 | 0.92989 | 0.9907 | 1.0084 | 0.98334 | Ce_ppm_m140 | 0.92204 | 0.97697 | 0.97697 | 0.94558 | 0.97108 | 0.9907 | 0.96389
Pr_ppm_m141 | 1.0645 | 0.92217 | 0.91134 | 0.93146 | 0.95737 | Pr_ppm_m141 | 0.89123 | 1.0429 | 0.99644 | 0.91289 | 0.92836 | 0.94074 | 0.95209
Nd_ppm_m146 | 1.0424 | 0.97746 | 0.89811 | 0.88729 | 0.95131 | Nd_ppm_m146 | 0.89811 | 0.95221 | 0.96303 | 0.90893 | 0.92696 | 0.92335 | 0.92876
Sm_ppm_m147 | 0.91954 | 1.0221 | 1.0221 | 1.1028 | 1.0166 | Sm_ppm_m147 | 0.95682 | 0.96303 | 0.94284 | 0.95061 | 0.93352 | 0.92265 | 0.94491
Eu_ppm_m153 | 1.0266 | 0.91714 | 0.92236 | 1.136 | 1.0005 | Eu_ppm_m153 | 0.95883 | 0.87024 | 0.92757 | 0.87546 | 0.96404 | 0.94841 | 0.92409
Gd_ppm_m157 | 0.94654 | 0.90566 | 0.90881 | 0.90252 | 0.91588 | Gd_ppm_m157 | 0.93711 | 0.98742 | 0.95912 | 0.85849 | 0.96384 | 0.99843 | 0.95073
Tb_ppm_m159 | 0.88606 | 1.0275 | 1.0071 | 0.92574 | 0.9616 | Tb_ppm_m159 | 0.91251 | 1.0285 | 0.95829 | 0.96338 | 0.92269 | 0.9471 | 0.95541
Dy_ppm_m163 | 0.88057 | 1.0005 | 1.0169 | 0.93971 | 0.95942 | Dy_ppm_m163 | 0.94299 | 0.98078 | 1.0071 | 0.98735 | 0.91506 | 0.97256 | 0.96764
Ho_ppm_m165 | 0.96955 | 0.91587 | 0.82532 | 0.96154 | 0.91807 | Ho_ppm_m165 | 0.90064 | 1.0393 | 0.97756 | 0.95433 | 0.9375 | 0.97756 | 0.96448
Er_ppm_m166 | 1.1748 | 0.83712 | 0.87401 | 0.94211 | 0.95701 | Er_ppm_m166 | 0.98184 | 1.0187 | 0.94211 | 0.90522 | 0.97616 | 1.0216 | 0.97427
Tm_ppm_m169 | 0.98361 | 0.98156 | 1.0102 | 1.0758 | 1.0128 | Tm_ppm_m169 | 0.8832 | 0.96926 | 0.9877 | 0.89139 | 0.90984 | 1.0328 | 0.9457
Yb_ppm_m172 | 0.89844 | 0.96755 | 1.0457 | 1.0457 | 0.98933 | Yb_ppm_m172 | 0.94952 | 0.9976 | 1.0306 | 1.0276 | 0.99159 | 0.90745 | 0.98407
Lu_ppm_m175 | 0.89691 | 0.90103 | 0.81649 | 1.134 | 0.93711 | Lu_ppm_m175 | 0.94433 | 0.94639 | 0.98144 | 0.87216 | 0.93196 | 0.97526 | 0.94192
Hf_ppm_m177 | 0.94895 | 1.0294 | 0.86846 | 0.98284 | 0.95742 | Hf_ppm_m177 | 0.9426 | 1.0655 | 1.0782 | 0.86211 | 0.95531 | 0.92989 | 0.97225
Ta_ppm_m181 | 0.9 | 1 | 0.94722 | 0.94444 | 0.94792 | Ta_ppm_m181 | 0.96806 | 0.9375 | 0.99722 | 0.92361 | 1.0167 | 0.94444 | 0.96458
W_ppm_m182 | 0.9405 | 0.762 | 0.90403 | 0.91363 | 0.88004 | W_ppm_m182 | 1.0576 | 0.97889 | 1.0038 | 0.93858 | 1.0883 | 1.1324 | 1.0333
Tl_ppm_m205 | 1.0448 | 1.204 | 0.96517 | 1.0249 | 1.0597 | Tl_ppm_m205 | 1.1493 | 1.0945 | 1.1592 | 1.0846 | 1.3383 | 1.1592 | 1.1642
Pb_ppm_m208 | 1.0844 | 0.99408 | 0.9278 | 1.1045 | 1.0277 | Pb_ppm_m208 | 1.0503 | 1.0202 | 1.0583 | 1.0041 | 0.9991 | 1.0433 | 1.0292
Bi_ppm_m209 | 0.64602 | 0.72678 | 1.1507 | 1.2517 | 0.9438 | Bi_ppm_m209 | 1.3526 | 0.98923 | 0.82772 | 0.70659 | 1.0902 | 0.64602 | 0.93539
Th_ppm_m232 | 0.91207 | 1.0276 | 1.0121 | 0.98448 | 0.98405 | Th_ppm_m232 | 0.95172 | 0.9569 | 0.96724 | 0.95172 | 0.96034 | 0.92241 | 0.95172
U_ppm_m238 | 0.94186 | 0.9186 | 0.90698 | 0.97674 | 0.93605 | U_ppm_m238 | 0.93023 | 0.94767 | 0.91047 | 0.96512 | 0.93023 | 0.98256 | 0.94438
Th/U | 0.96837 | 1.1186 | 1.1159 | 1.0079 | 1.0513 | Th/U | 1.0231 | 1.0097 | 1.0624 | 0.98612 | 1.0324 | 0.93879 | 1.0078

## EPMA secondary standards - VG2
2019-08-01
Preferred values
| Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl
| 2.6498 | 13.63 | 0.2205 | 10.8534 | 0.1987 | 1.891 | 50.26 | 6.839 | 11.744 | 0.225 | 0.2954 | 0.0304
MonMorning | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6781 | 13.5066 | 0.1999 | 10.9812 | 0.2141 | 1.8619 | 50.0813 | 6.6852 | 11.4792 | 0.1535 | 0.2914 | 0.0313 | 98.1638 | 5874 | 27134 | 15 |  |  | VG2_crystal2_2 |  | 12.6005 | 1 | 2019-08-12 11:53
2 / 1 . | 2.7841 | 13.5924 | 0.1671 | 10.8458 | 0.1665 | 1.8539 | 50.146 | 6.6356 | 11.5655 | 0.2148 | 0.2985 | 0.027 | 98.2972 | 5862 | 27121 | 15 |  |  | VG2_crystal2_3 |  | 12.6188 | 2 | 2019-08-12 12:02
3 / 1 . | 2.6231 | 13.3696 | 0.2092 | 10.7524 | 0.1793 | 1.8698 | 50.2562 | 6.578 | 11.704 | 0.1868 | 0.2972 | 0.0292 | 98.0548 | 5866 | 27107 | 15 |  |  | VG2_crystal2_4 |  | 12.6032 | 3 | 2019-08-12 12:10
Recovery | 101.7096 | 98.9694 | 87.1051 | 100.059 | 93.9272 | 98.4594 | 99.8034 | 96.9869 | 98.6282 | 82.237 | 100.1016 | 95.943 | #DIV/0!
Mon Evening (1) befpre chain aq
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6514 | 13.3235 | 0.1812 | 10.8807 | 0.1794 | 1.861 | 49.5916 | 6.5073 | 11.8485 | 0.2576 | 0.2968 | 0.0296 | 97.6087 | 5888 | 27134 | 15 |  |  | VG2_overNtMon1_1 |  | 12.5856 | 1 | 2019-08-12 17:21
2 / 1 . | 2.6314 | 13.1681 | 0.1519 | 10.857 | 0.1969 | 1.8316 | 49.168 | 6.6185 | 11.9232 | 0.1735 | 0.3046 | 0.0336 | 97.0582 | 5877 | 27157 | 15 |  |  | VG2_overNtMon1_2 |  | 12.524 | 2 | 2019-08-12 17:29
3 / 1 . | 2.6727 | 13.2394 | 0.233 | 10.9197 | 0.1883 | 1.8689 | 49.557 | 6.4535 | 11.5428 | 0.1541 | 0.3006 | 0.0276 | 97.1576 | 5874 | 27177 | 15 |  |  | VG2_overNtMon1_3 |  | 12.4955 | 3 | 2019-08-12 17:38
Recovery | 100.0767 | 97.1656 | 85.5782 | 100.2985 | 94.7157 | 98.0345 | 98.3662 | 95.4296 | 100.2342 | 86.6963 | 101.7829 | 99.5614 | #DIV/0!
Mon night mid | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.663 | 13.0671 | 0.236 | 10.6564 | 0.1594 | 1.8485 | 49.6824 | 6.5575 | 11.4626 | 0.2063 | 0.2846 | 0.03 | 96.8538 | 5852 | 27171 | 15 |  |  | VG2_overNtMon1_4 |  | 12.4408 | 1 | 2019-08-12 23:01
2 / 1 . | 2.6456 | 12.9934 | 0.2178 | 10.8051 | 0.2199 | 1.8664 | 49.3033 | 6.5483 | 11.5458 | 0.1691 | 0.2893 | 0.0318 | 96.6357 | 5850 | 27186 | 15 |  |  | VG2_overNtMon1_5 |  | 12.4365 | 2 | 2019-08-12 23:09
3 / 1 . | 2.7198 | 13.1129 | 0.1586 | 11.0921 | 0.1966 | 1.8764 | 49.8526 | 6.8022 | 11.6339 | 0.1987 | 0.2968 | 0.0358 | 97.9764 | 5841 | 27210 | 15 |  |  | VG2_overNtMon1_6 |  | 12.6085 | 3 | 2019-08-12 23:17
Recovery | 100.9938 | 95.8019 | 92.5775 | 99.9797 | 96.6113 | 98.5598 | 98.7122 | 97.0317 | 98.3262 | 85.0519 | 98.251 | 107.0175 | #DIV/0!
Mon night end | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
4 / 1 . | 2.6553 | 13.0899 | 0.1914 | 10.7279 | 0.2008 | 1.8704 | 49.2538 | 6.5394 | 11.8459 | 0.1229 | 0.2915 | 0.0345 | 96.8238 | 5873 | 27223 | 15 |  |  | VG2 |  | 12.4801 | 4 | 2019-08-13 06:43
5 / 1 . | 2.5707 | 13.0897 | 0.2352 | 10.7457 | 0.1754 | 1.8636 | 49.8746 | 6.5663 | 11.5452 | 0.1856 | 0.2954 | 0.0337 | 97.1809 | 5825 | 27166 | 15 |  |  | VG2 |  | 12.4911 | 5 | 2019-08-13 06:51
6 / 1 . | 2.6251 | 13.0778 | 0.2558 | 10.6176 | 0.1913 | 1.8802 | 49.5899 | 6.6576 | 11.3335 | 0.212 | 0.2987 | 0.03 | 96.7694 | 5796 | 27184 | 15 |  |  | VG2 |  | 12.4194 | 6 | 2019-08-13 07:00
Recovery | 98.7634 | 96.0073 | 103.1595 | 98.5596 | 95.2021 | 98.9635 | 98.6326 | 96.3265 | 98.5598 | 77.1111 | 99.9323 | 107.6754 | #DIV/0!
tue midday | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.5763 | 13.2624 | 0.174 | 10.9543 | 0.1899 | 1.8472 | 50.2355 | 6.5713 | 11.638 | 0.2061 | 0.2995 | 0.0346 | 97.9893 | 5856 | 27163 | 15 |  |  | VG2_overNtMon1_4 |  | 12.6026 | 1 | 2019-08-13 11:53
2 / 1 . | 2.6534 | 13.2268 | 0.265 | 10.7708 | 0.2015 | 1.8782 | 49.9279 | 6.6507 | 11.5581 | 0.1868 | 0.2903 | 0.0361 | 97.6456 | 5846 | 27203 | 15 |  |  | VG2_overNtMon1_5 |  | 12.5459 | 2 | 2019-08-13 12:01
3 / 1 . | 2.6453 | 13.2134 | 0.2081 | 10.6956 | 0.1511 | 1.8594 | 50.4444 | 6.6132 | 11.7692 | 0.24 | 0.3068 | 0.0334 | 98.1797 | 5841 | 27223 | 15 |  |  | VG2_overNtMon1_6 |  | 12.6242 | 3 | 2019-08-13 12:10
Recovery | 99.0641 | 97.0961 | 97.8231 | 99.5716 | 91.0082 | 98.4453 | 99.8858 | 96.6769 | 99.243 | 93.763 | 101.1735 | 114.1447 | #DIV/0!
wednightstart | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
4 / 1 . | 2.5844 | 13.2077 | 0.216 | 11.0104 | 0.2061 | 1.8205 | 48.6363 | 6.4464 | 11.5693 | 0.2243 | 0.2975 | 0.028 | 96.2469 | 5871 | 27223 | 4 |  |  | VG2 |  | 12.4115 | 4 | 2019-08-14 20:34
5 / 1 . | 2.5976 | 13.145 | 0.2188 | 10.8887 | 0.2108 | 1.8448 | 48.7434 | 6.5689 | 11.259 | 0.2067 | 0.2992 | 0.0311 | 96.014 | 5831 | 27166 | 5 |  |  | VG2 |  | 12.3442 | 5 | 2019-08-14 20:42
6 / 1 . | 2.617 | 13.2848 | 0.2214 | 10.939 | 0.166 | 1.8377 | 50.1763 | 6.7327 | 11.4736 | 0.1736 | 0.3015 | 0.0289 | 97.9526 | 5865 | 27176 | 9 |  |  | VG2 |  | 12.5728 | 6 | 2019-08-14 20:50
Recovery | 98.108 | 96.9369 | 99.1988 | 100.8535 | 97.7856 | 97.0033 | 97.8618 | 96.2519 | 97.3601 | 89.5704 | 101.3541 | 96.4912 | #DIV/0!
wed night mid | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6613 | 13.0077 | 0.1832 | 10.9595 | 0.2148 | 1.8411 | 48.5108 | 6.5393 | 11.561 | 0.2546 | 0.2946 | 0.0336 | 96.0614 | 5851 | 27166 | 6 |  |  | VG2_overNtMon1_4 |  | 12.3922 | 1 | 2019-08-15 02:17
2 / 1 . | 2.6071 | 13.0515 | 0.2566 | 11.0194 | 0.1828 | 1.8765 | 48.988 | 6.5557 | 11.4084 | 0.1953 | 0.3019 | 0.0283 | 96.4716 | 5850 | 27222 | 7 |  |  | VG2_overNtMon1_5 |  | 12.4168 | 2 | 2019-08-15 02:26
3 / 1 . | 2.5812 | 13.1058 | 0.2608 | 11.0322 | 0.1766 | 1.8661 | 49.8242 | 6.597 | 11.4035 | 0.2121 | 0.2883 | 0.0358 | 97.3836 | 5899 | 27210 | 6 |  |  | VG2_overNtMon1_6 |  | 12.5165 | 3 | 2019-08-15 02:34
Recovery | 98.7446 | 95.7814 | 105.9108 | 101.3848 | 96.3261 | 98.4259 | 97.7073 | 95.9789 | 97.5616 | 98.0741 | 99.842 | 107.1272 | #DIV/0!
wednight end
4 / 1 . | 2.76 | 13.1657 | 0.1836 | 10.9104 | 0.1734 | 1.8788 | 48.7254 | 6.5117 | 11.707 | 0.2135 | 0.3009 | 0.0344 | 96.5648 | 6010 | 27246 | 6 |  |  | VG2 |  | 12.4547 | 4 | 2019-08-15 08:35
5 / 1 . | 2.7046 | 13.1578 | 0.1998 | 10.7236 | 0.1803 | 1.8592 | 49.5791 | 6.5358 | 11.3463 | 0.175 | 0.3007 | 0.0277 | 96.79 | 5872 | 27122 | 11 |  |  | VG2 |  | 12.4231 | 5 | 2019-08-15 08:43
6 / 1 . | 2.733 | 13.2339 | 0.2246 | 10.8573 | 0.1953 | 1.8396 | 49.5108 | 6.6174 | 11.772 | 0.1623 | 0.3061 | 0.0307 | 97.483 | 5881 | 27144 | 9 |  |  | VG2 |  | 12.5516 | 6 | 2019-08-15 08:52
Recovery | 103.1222 | 96.741 | 91.9123 | 99.7884 | 92.0986 | 98.3184 | 98.0338 | 95.8469 | 98.8457 | 81.6 | 102.4261 | 101.7544 | #DIV/0!
Thur middday | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6184 | 13.3145 | 0.1668 | 10.7892 | 0.2138 | 1.8529 | 49.7597 | 6.8333 | 11.8684 | 0.2429 | 0.2953 | 0.0332 | 97.9885 | 5843 | 27171 | 12 |  |  | VG2_overNtMon1_4 |  | 12.6229 | 1 | 2019-08-15 14:42
2 / 1 . | 2.5642 | 13.1909 | 0.168 | 11.0884 | 0.1951 | 1.8281 | 50.1712 | 6.7009 | 11.716 | 0.2616 | 0.2975 | 0.0299 | 98.212 | 5833 | 27222 | 9 |  |  | VG2_overNtMon1_5 |  | 12.6475 | 2 | 2019-08-15 14:50
3 / 1 . | 2.751 | 13.3741 | 0.2415 | 10.9795 | 0.1947 | 1.8456 | 49.7977 | 6.6217 | 11.8049 | 0.2517 | 0.2955 | 0.0314 | 98.1893 | 5899 | 27221 | 11 |  |  | VG2_overNtMon1_6 |  | 12.6478 | 3 | 2019-08-15 14:59
Recovery | 99.8012 | 97.5287 | 87.1202 | 100.9118 | 101.2582 | 97.4194 | 99.3027 | 98.24 | 100.4465 | 112.0296 | 100.237 | 103.6184 | #DIV/0!
Thur night
4 / 1 . | 2.6844 | 13.1105 | 0.1955 | 10.9271 | 0.1959 | 1.847 | 49.1396 | 6.7767 | 11.5256 | 0.2748 | 0.3057 | 0.0355 | 97.0183 | 5869 | 27233 | 8 |  |  | VG2 |  | 12.49 | 4 | 2019-08-15 18:37
5 / 1 . | 2.4494 | 13.314 | 0.2107 | 10.9814 | 0.2126 | 1.831 | 49.386 | 6.7361 | 11.4823 | 0.1475 | 0.2977 | 0.032 | 97.0807 | 5822 | 27155 | 14 |  |  | VG2 |  | 12.4829 | 5 | 2019-08-15 18:45
6 / 1 . | 2.6749 | 13.0351 | 0.2038 | 10.8248 | 0.1876 | 1.8611 | 49.449 | 6.601 | 11.4763 | 0.1694 | 0.2912 | 0.0344 | 96.8086 | 5893 | 27176 | 9 |  |  | VG2 |  | 12.4459 | 6 | 2019-08-15 18:54
Recovery | 98.2301 | 96.5018 | 92.2147 | 100.5316 | 100 | 97.6397 | 98.1394 | 98.0348 | 97.8775 | 87.6593 | 100.9479 | 111.7325 | #DIV/0!
Overall recovery | 99.8614 | 96.853 | 94.26 | 100.1939 | 95.8933 | 98.1269 | 98.6445 | 96.6804 | 98.7083 | 89.3793 | 100.6048 | 104.5066 | #DIV/0!
Overall precision | 2.5331 | 1.0537 | 14.9382 | 1.143 | 9.055 | 0.87543 | 1.0261 | 1.4269 | 1.4616 | 18.6179 | 1.7985 | 8.5979
2020-01-01
| Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl
Preferred | 2.667 | 13.574 | 0.202 | 50.249 | 6.817 | 0.189 | 10.822 | 1.898 | 11.736 | 0.198 | 0.299 | 0.031
Mon start | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
3 / 1 . | 2.9137 | 13.1875 | 0.2038 | 49.6649 | 6.6527 | 0.185 | 10.7366 | 1.8442 | 11.2855 | 0.1847 | 0.2988 | 0.0286 | 97.1857 | 6062 | 26857 | -62 |  |  | VG2_crystal2_1 |  | 12.4586 | 3 | 2020-01-28 14:27
4 / 1 . | 2.8262 | 13.4812 | 0.2872 | 49.0952 | 6.768 | 0.2068 | 10.86 | 1.8514 | 11.9122 | 0.1961 | 0.3087 | 0.0275 | 97.8205 | 6069 | 26846 | -62 |  |  | VG2_crystal2_2 |  | 12.6072 | 4 | 2020-01-28 14:36
5 / 1 . | 2.8879 | 13.2219 | 0.1757 | 49.3043 | 6.7733 | 0.2495 | 10.7458 | 1.858 | 11.7482 | 0.1294 | 0.3006 | 0.0292 | 97.4238 | 6036 | 26899 | -62 |  |  | VG2_crystal2_3 |  | 12.5353 | 5 | 2020-01-28 14:44
av/Pref | 1.0783 | 0.97958 | 1.1002 | 0.9822 | 0.98743 | 1.131 | 0.99619 | 0.97534 | 0.99256 | 0.85892 | 1.0124 | 0.9172 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
[13 empty] | 0 | . | . | . |  |  | VG2_Crystal2_2 |  | 0 | 5 | 2020-01-28 13:00
Mon night standards | [12 empty] | 0 | . | . | . |  |  | VG2_Crystal2_3 |  | 0 | 6 | 2020-01-28 13:01
4 / 1 . | 2.7611 | 12.8831 | 0.1783 | 48.8923 | 6.6553 | 0.2277 | 10.6983 | 1.8726 | 11.7421 | 0.2632 | 0.2996 | 0.0333 | 96.507 | 6077 | 26861 | -62 |  |  | VG2_crystal2_1 |  | 12.4482 | 4 | 2020-01-29 02:40
5 / 1 . | 2.7743 | 12.9146 | 0.1566 | 49.4688 | 6.6077 | 0.1622 | 10.6701 | 1.8674 | 11.4398 | 0.2075 | 0.2894 | 0.038 | 96.5963 | 6079 | 26829 | -62 |  |  | VG2_crystal2_2 |  | 12.4123 | 5 | 2020-01-29 02:49
6 / 1 . | 2.6219 | 12.8525 | 0.1874 | 49.1469 | 6.5952 | 0.2095 | 10.7165 | 1.8905 | 12.0647 | 0.1631 | 0.3059 | 0.0346 | 96.7886 | 5998 | 26884 | -62 |  |  | VG2_crystal2_3 |  | 12.5065 | 6 | 2020-01-29 02:57
av/Pref | 101.9535 | 94.9123 | 86.1881 | 97.8514 | 97.1014 | 105.7143 | 98.8262 | 98.8848 | 100.1096 | 106.7003 | 99.7659 | 113.871 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Tue lunchtime standards
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.9865 | 13.3267 | 0.1812 | 49.0296 | 6.6449 | 0.2227 | 10.7409 | 1.8646 | 11.5644 | 0.1868 | 0.3065 | 0.031 | 97.0858 | 5918 | 26830 | -58 |  |  | VG2_1 |  | 12.4826 | 1 | 2020-01-29 13:16
2 / 1 . | 2.7739 | 13.1622 | 0.1773 | 48.9 | 6.5553 | 0.1766 | 10.7572 | 1.8715 | 11.5239 | 0.2072 | 0.3078 | 0.0311 | 96.444 | 6049 | 26892 | -59 |  |  | VG2_2 |  | 12.4111 | 2 | 2020-01-29 13:25
3 / 1 . | 2.9015 | 12.8378 | 0.249 | 48.7899 | 6.6399 | 0.2112 | 10.782 | 1.8777 | 11.7059 | 0.1774 | 0.2983 | 0.0322 | 96.5027 | 6056 | 26884 | -62 |  |  | VG2_3 |  | 12.4382 | 3 | 2020-01-29 13:33
av/Pref | 108.2602 | 96.5736 | 100.2475 | 97.3283 | 97.0129 | 107.672 | 99.4274 | 98.5915 | 98.8247 | 96.1953 | 101.7391 | 101.3978 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
2020-02-01
| Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl
| 2.667 | 13.574 | 0.202 | 50.249 | 6.817 | 0.189 | 10.822 | 1.898 | 11.736 | 0.198 | 0.299 | 0.031
Mon start | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
3 / 1 . | 2.9137 | 13.1875 | 0.2038 | 49.6649 | 6.6527 | 0.185 | 10.7366 | 1.8442 | 11.2855 | 0.1847 | 0.2988 | 0.0286 | 97.1857 | 6062 | 26857 | -62 |  |  | VG2_crystal2_1 |  | 12.4586 | 3 | 2020-01-28 14:27
4 / 1 . | 2.8262 | 13.4812 | 0.2872 | 49.0952 | 6.768 | 0.2068 | 10.86 | 1.8514 | 11.9122 | 0.1961 | 0.3087 | 0.0275 | 97.8205 | 6069 | 26846 | -62 |  |  | VG2_crystal2_2 |  | 12.6072 | 4 | 2020-01-28 14:36
5 / 1 . | 2.8879 | 13.2219 | 0.1757 | 49.3043 | 6.7733 | 0.2495 | 10.7458 | 1.858 | 11.7482 | 0.1294 | 0.3006 | 0.0292 | 97.4238 | 6036 | 26899 | -62 |  |  | VG2_crystal2_3 |  | 12.5353 | 5 | 2020-01-28 14:44
av/Pref | 1.0783 | 0.97958 | 1.1002 | 0.9822 | 0.98743 | 1.131 | 0.99619 | 0.97534 | 0.99256 | 0.85892 | 1.0124 | 0.9172 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
[13 empty] | 0 | . | . | . |  |  | VG2_Crystal2_2 |  | 0 | 5 | 2020-01-28 13:00
Mon night standards | [12 empty] | 0 | . | . | . |  |  | VG2_Crystal2_3 |  | 0 | 6 | 2020-01-28 13:01
4 / 1 . | 2.7611 | 12.8831 | 0.1783 | 48.8923 | 6.6553 | 0.2277 | 10.6983 | 1.8726 | 11.7421 | 0.2632 | 0.2996 | 0.0333 | 96.507 | 6077 | 26861 | -62 |  |  | VG2_crystal2_1 |  | 12.4482 | 4 | 2020-01-29 02:40
5 / 1 . | 2.7743 | 12.9146 | 0.1566 | 49.4688 | 6.6077 | 0.1622 | 10.6701 | 1.8674 | 11.4398 | 0.2075 | 0.2894 | 0.038 | 96.5963 | 6079 | 26829 | -62 |  |  | VG2_crystal2_2 |  | 12.4123 | 5 | 2020-01-29 02:49
6 / 1 . | 2.6219 | 12.8525 | 0.1874 | 49.1469 | 6.5952 | 0.2095 | 10.7165 | 1.8905 | 12.0647 | 0.1631 | 0.3059 | 0.0346 | 96.7886 | 5998 | 26884 | -62 |  |  | VG2_crystal2_3 |  | 12.5065 | 6 | 2020-01-29 02:57
4 / 1 . | 101.9535 | 94.9123 | 86.1881 | 97.8514 | 97.1014 | 105.7143 | 98.8262 | 98.8848 | 100.1096 | 106.7003 | 99.7659 | 113.871 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
wed lunchtime standards
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.9865 | 13.3267 | 0.1812 | 49.0296 | 6.6449 | 0.2227 | 10.7409 | 1.8646 | 11.5644 | 0.1868 | 0.3065 | 0.031 | 97.0858 | 5918 | 26830 | -58 |  |  | VG2_1 |  | 12.4826 | 1 | 2020-01-29 13:16
2 / 1 . | 2.7739 | 13.1622 | 0.1773 | 48.9 | 6.5553 | 0.1766 | 10.7572 | 1.8715 | 11.5239 | 0.2072 | 0.3078 | 0.0311 | 96.444 | 6049 | 26892 | -59 |  |  | VG2_2 |  | 12.4111 | 2 | 2020-01-29 13:25
3 / 1 . | 2.9015 | 12.8378 | 0.249 | 48.7899 | 6.6399 | 0.2112 | 10.782 | 1.8777 | 11.7059 | 0.1774 | 0.2983 | 0.0322 | 96.5027 | 6056 | 26884 | -62 |  |  | VG2_3 |  | 12.4382 | 3 | 2020-01-29 13:33
av/Pref | 108.2602 | 96.5736 | 100.2475 | 97.3283 | 97.0129 | 107.672 | 99.4274 | 98.5915 | 98.8247 | 96.1953 | 101.7391 | 101.3978 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Wed evening end
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
4 / 1 . | 2.7754 | 13.788 | 0.174 | 49.7382 | 6.7519 | 0.201 | 10.8038 | 1.8872 | 11.4656 | 0.1374 | 0.3015 | 0.0376 | 98.0615 | 6077 | 26842 | -56 |  |  | VG2_crystal2_1 |  | 12.5757 | 4 | 2020-01-29 18:38
5 / 1 . | 2.727 | 13.7311 | 0.1708 | 49.499 | 6.6316 | 0.2134 | 10.6625 | 1.8275 | 11.3436 | 0.2649 | 0.2989 | 0.034 | 97.4043 | 6095 | 26844 | -56 |  |  | VG2_crystal2_2 |  | 12.4942 | 5 | 2020-01-29 18:46
6 / 1 . | 2.8558 | 13.5358 | 0.2263 | 49.2607 | 6.6721 | 0.2009 | 10.901 | 1.8632 | 11.1184 | 0.1522 | 0.3011 | 0.0328 | 97.1205 | 6011 | 26880 | -56 |  |  | VG2_crystal2_3 |  | 12.4411 | 6 | 2020-01-29 18:55
av/Pref | 104.4644 | 100.8175 | 94.2409 | 98.508 | 98.0666 | 108.5185 | 99.696 | 97.961 | 96.3633 | 93.3502 | 100.5017 | 112.2581 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Thur cofefe
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.8589 | 13.6438 | 0.2212 | 49.5043 | 6.6524 | 0.1966 | 10.7825 | 1.8685 | 11.8277 | 0.2011 | 0.3047 | 0.032 | 98.0935 | 5902 | 26893 | -58 |  |  | VG2_1 |  | 12.6237 | 1 | 2020-01-30 11:06
2 / 1 . | 2.744 | 13.423 | 0.1839 | 48.9975 | 6.4875 | 0.2064 | 10.8919 | 1.8562 | 11.7913 | 0.2412 | 0.3035 | 0.031 | 97.1574 | 6052 | 26904 | -59 |  |  | VG2_2 |  | 12.5304 | 2 | 2020-01-30 11:14
3 / 1 . | 2.7437 | 13.2908 | 0.2135 | 49.3561 | 6.8653 | 0.2073 | 10.8381 | 1.8694 | 11.8604 | 0.2243 | 0.2865 | 0.0321 | 97.7875 | 6078 | 26854 | -62 |  |  | VG2_3 |  | 12.6008 | 3 | 2020-01-30 11:23
av/Pref | 104.3195 | 99.1052 | 102.0792 | 98.0835 | 97.8202 | 107.6367 | 100.1432 | 98.2455 | 100.7708 | 112.2222 | 99.7436 | 102.2581 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Thur night end of glass run
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.7646 | 13.4393 | 0.2056 | 49.1532 | 6.7767 | 0.1796 | 10.7981 | 1.8796 | 11.8183 | 0.2438 | 0.303 | 0.0335 | 97.5953 | 5805 | 26893 | -58 |  |  | VG2_1 |  | 12.5741 | 1 | 2020-01-30 22:42
2 / 1 . | 2.7735 | 13.5817 | 0.1908 | 49.3874 | 6.7603 | 0.2067 | 10.6605 | 1.8388 | 11.6457 | 0.1844 | 0.3025 | 0.0326 | 97.5648 | 5802 | 26874 | -58 |  |  | VG2_2 |  | 12.5365 | 2 | 2020-01-30 22:51
3 / 1 . | 2.8439 | 13.6206 | 0.2264 | 49.4265 | 6.7832 | 0.2306 | 10.8144 | 1.8699 | 11.4855 | 0.194 | 0.2975 | 0.0352 | 97.8278 | 5806 | 26862 | -58 |  |  | VG2_3 |  | 12.5598 | 3 | 2020-01-30 22:59
av/Pref | 104.7619 | 99.8026 | 102.7723 | 98.1559 | 99.3604 | 108.8007 | 99.4055 | 98.1437 | 99.2658 | 104.7475 | 100.6689 | 108.9247 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
fri night start
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
4 / 1 . | 2.8987 | 13.4996 | 0.1378 | 49.0294 | 6.7455 | 0.2046 | 10.9834 | 1.848 | 11.5626 | 0.2021 | 0.2951 | 0.0332 | 97.4399 | 6073 | 26840 | -56 |  |  | VG2_crystal2_1 |  | 12.5339 | 4 | 2020-01-31 19:00
5 / 1 . | 2.8347 | 13.6304 | 0.2907 | 49.0931 | 6.4918 | 0.165 | 10.9052 | 1.8663 | 11.0103 | 0.187 | 0.2907 | 0.0311 | 96.7962 | 6107 | 26820 | -56 |  |  | VG2_crystal2_2 |  | 12.3963 | 5 | 2020-01-31 19:08
6 / 1 . | 2.698 | 13.4974 | 0.1925 | 49.0117 | 6.7102 | 0.1993 | 10.7965 | 1.8586 | 11.5742 | 0.193 | 0.2972 | 0.033 | 97.0617 | 6031 | 26800 | -56 |  |  | VG2_crystal2_3 |  | 12.4843 | 6 | 2020-01-31 19:17
av/Pref | 105.3793 | 99.7677 | 102.4752 | 97.6034 | 97.538 | 100.3351 | 100.6749 | 97.8732 | 96.9868 | 97.9966 | 98.4392 | 104.6237 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Fri night end | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.8528 | 13.7675 | 0.1894 | 49.4845 | 6.6641 | 0.1755 | 10.7757 | 1.8876 | 11.7112 | 0.2472 | 0.2977 | 0.0298 | 98.083 | 5882 | 26779 | -54 |  |  | VG2_1 |  | 12.6131 | 1 | 2020-02-01 01:28
2 / 1 . | 2.8023 | 13.7432 | 0.214 | 49.1733 | 6.6874 | 0.2139 | 10.9001 | 1.8579 | 11.6304 | 0.217 | 0.2955 | 0.0329 | 97.768 | 5890 | 26737 | -54 |  |  | VG2_2 |  | 12.5753 | 2 | 2020-02-01 01:37
3 / 1 . | 2.8375 | 13.7732 | 0.1751 | 49.3143 | 6.6965 | 0.1877 | 10.7719 | 1.868 | 11.3063 | 0.191 | 0.3029 | 0.0327 | 97.4572 | 5898 | 26737 | -54 |  |  | VG2_3 |  | 12.494 | 3 | 2020-02-01 01:45
av/Pref | 106.1442 | 101.3798 | 95.462 | 98.1592 | 98.0294 | 101.7813 | 99.9436 | 98.5862 | 98.4092 | 110.303 | 99.8997 | 102.5806 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Sat afternoon | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6986 | 13.4888 | 0.2235 | 49.483 | 6.7621 | 0.1796 | 10.9239 | 1.8579 | 11.7716 | 0.1749 | 0.3028 | 0.0358 | 97.9024 | 5884 | 26769 | -54 |  |  | VG2_1 |  | 12.6015 | 1 | 2020-02-01 14:19
2 / 1 . | 2.7279 | 13.3675 | 0.2101 | 48.5713 | 6.6385 | 0.1988 | 10.9671 | 1.8908 | 11.329 | 0.2283 | 0.3138 | 0.0309 | 96.474 | 5904 | 26730 | -54 |  |  | VG2_2 |  | 12.409 | 2 | 2020-02-01 14:27
3 / 1 . | 2.7659 | 13.5203 | 0.1775 | 48.9363 | 6.5681 | 0.1836 | 10.8102 | 1.8334 | 11.4685 | 0.1774 | 0.304 | 0.0319 | 96.7771 | 5905 | 26681 | -54 |  |  | VG2_3 |  | 12.4386 | 3 | 2020-02-01 14:35
av/Pref | 102.3922 | 99.1518 | 100.8416 | 97.5081 | 97.6417 | 99.1182 | 100.7245 | 98.0348 | 98.1854 | 97.7441 | 102.631 | 106.0215 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.7581 | 13.3331 | 0.1857 | 49.2364 | 6.6206 | 0.2312 | 10.9391 | 1.8765 | 11.0818 | 0.1549 | 0.3004 | 0.036 | 96.7538 | 5897 | 26779 | -54 |  |  | VG2_1 |  | 12.4037 | 1 | 2020-02-02 03:40
2 / 1 . | 2.8763 | 13.2636 | 0.2883 | 49.15 | 6.7382 | 0.2 | 10.7508 | 1.8597 | 11.7014 | 0.215 | 0.3044 | 0.0349 | 97.3826 | 5879 | 26725 | -54 |  |  | VG2_2 |  | 12.5326 | 2 | 2020-02-02 03:48
3 / 1 . | 2.7772 | 13.3055 | 0.2306 | 49.5455 | 6.7152 | 0.2131 | 10.8789 | 1.8367 | 11.8324 | 0.1742 | 0.3029 | 0.0363 | 97.8484 | 5921 | 26724 | -54 |  |  | VG2_3 |  | 12.6008 | 3 | 2020-02-02 03:57
av/Pref | 105.1319 | 97.9868 | 116.2706 | 98.1326 | 98.1566 | 113.6332 | 100.3166 | 97.8732 | 98.3174 | 91.5993 | 101.1929 | 115.2688 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Sat night 2
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.4569 | 13.6926 | 0.1908 | 49.9575 | 6.7822 | 0.1966 | 10.8083 | 1.8579 | 11.7283 | 0.2085 | 0.3052 | 0.0324 | 98.2173 | 5830 | 26819 | -54 |  |  | VG2_1 |  | 12.6293 | 1 | 2020-02-02 08:34
2 / 1 . | 2.6601 | 13.4554 | 0.1873 | 48.9496 | 6.479 | 0.1653 | 10.9138 | 1.8607 | 11.2141 | 0.2057 | 0.3165 | 0.0373 | 96.4451 | 5922 | 26725 | -54 |  |  | VG2_2 |  | 12.3851 | 2 | 2020-02-02 08:42
3 / 1 . | 2.7831 | 13.532 | 0.1905 | 49.265 | 6.6208 | 0.2267 | 10.8181 | 1.8705 | 11.4547 | 0.1492 | 0.3023 | 0.038 | 97.2506 | 5921 | 26669 | -54 |  |  | VG2_3 |  | 12.4906 | 3 | 2020-02-02 08:51
av/Pref | 98.7389 | 99.8969 | 93.8284 | 98.2919 | 97.2177 | 103.8095 | 100.2285 | 98.1577 | 97.6968 | 94.8485 | 103.01 | 115.8065 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Sun afternoon
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.7825 | 13.7921 | 0.216 | 49.3493 | 6.7402 | 0.2206 | 11.1508 | 1.8594 | 11.5222 | 0.2298 | 0.3038 | 0.032 | 98.1987 | 5862 | 26777 | -60 |  |  | VG2_1 |  | 12.626 | 1 | 2020-02-02 15:40
2 / 1 . | 2.622 | 13.0402 | 0.1797 | 49.2757 | 6.6355 | 0.2026 | 10.8064 | 1.8672 | 11.3693 | 0.2169 | 0.2931 | 0.0357 | 96.5444 | 5901 | 26733 | -65 |  |  | VG2_2 |  | 12.4107 | 2 | 2020-02-02 15:49
3 / 1 . | 2.7805 | 12.8891 | 0.2196 | 49.141 | 6.725 | 0.2028 | 10.9047 | 1.884 | 11.6075 | 0.1939 | 0.3053 | 0.034 | 96.8874 | 5877 | 26687 | -66 |  |  | VG2_3 |  | 12.4778 | 3 | 2020-02-02 15:57
av/Pref | 102.2997 | 97.5429 | 101.5347 | 98.0225 | 98.2871 | 110.4056 | 101.2194 | 98.5353 | 97.9863 | 107.8451 | 100.5797 | 109.3548 | #DIV/0! | . | . | . |  |  | VG2_Crystal2_1 |  | 0 | 4 | 2020-01-28 13:00
Overal average/preferred | 104.3066 | 98.7413 | 100.4964 | 97.9888 | 97.9146 | 106.7108 | 100.0188 | 98.2018 | 98.5143 | 99.9537 | 100.7841 | 107.0072 | #DIV/0!
std dev/average | 3.4526 | 2.1058 | 16.8572 | 0.56461 | 1.3357 | 9.8368 | 0.92835 | 0.82669 | 2.1031 | 16.3537 | 1.9924 | 7.7222 | 0.57519

## EPMA secondary standards - A99
2019-08-01
| Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl
| 2.677 | 12.157 | 0.455 | 9.032 | 0.861 | 4.081 | 50.888 | 5.024 | 13.383 | 0.241 | 0.03 | 0.023
Start of day mon | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point#
1 / 1 . | 2.5403 | 11.8763 | 0.4613 | 8.9618 | 0.8798 | 4.0955 | 49.8287 | 4.9077 | 12.9872 | 0.2414 | 0.0252 | 0.0192 |  | 96.8244 | -7637 | 22951 | 41 |  |  | A99_Crystal3_run1 |  | 12.691 | 1
2 / 1 . | 2.7385 | 11.9097 | 0.3771 | 9.0153 | 0.9277 | 4.1028 | 49.7421 | 4.8813 | 13.4946 | 0.2779 | 0.0268 | 0.0182 |  | 97.512 | -7610 | 22968 | 41 |  |  | A99_Crystal3_run2 |  | 12.831 | 2
3 / 1 . | 2.7496 | 11.7871 | 0.4752 | 8.8814 | 0.8646 | 4.0764 | 49.541 | 4.9004 | 13.076 | 0.2461 | 0.0243 | 0.0241 |  | 96.6462 | -7559 | 22968 | 41 |  |  | A99_Crystal3_run3 |  | 12.6748 | 3
recovery | 99.9676 | 97.538 | 96.2344 | 99.1235 | 103.4495 | 100.2589 | 97.6732 | 97.4615 | 98.5275 | 105.8645 | 84.7778 | 89.1304
Mon night mid | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.4811 | 11.9436 | 0.4659 | 9.0926 | 0.8872 | 4.063 | 51.0583 | 4.8457 | 13.1549 | 0.2032 | 0.0282 | 0.0258 | 98.2493 | -7596 | 23004 | 49 |  |  | A99_Crystal3_run4 |  | 12.8669 | 1 | 2019-08-13 01:57
2 / 1 . | 2.5987 | 11.9005 | 0.4442 | 9.0743 | 0.8904 | 4.0695 | 50.0508 | 4.8903 | 13.0892 | 0.2008 | 0.0297 | 0.0205 | 97.2587 | -7593 | 23023 | 49 |  |  | A99_Crystal3_run5 |  | 12.7506 | 2 | 2019-08-13 02:05
3 / 1 . | 2.7058 | 11.9739 | 0.5461 | 9.057 | 0.8895 | 4.1264 | 50.6177 | 4.9997 | 13.3291 | 0.1868 | 0.0283 | 0.0186 | 98.4789 | -7655 | 23023 | 49 |  |  | A99_Crystal3_run6 |  | 12.9089 | 3 | 2019-08-13 02:14
recovery | 96.9443 | 98.2095 | 106.6813 | 100.472 | 103.2559 | 100.1299 | 99.3861 | 97.7687 | 98.5658 | 81.7151 | 95.7778 | 94.058
final standards Monday night | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.7625 | 11.942 | 0.4153 | 9.2036 | 0.8677 | 4.1048 | 50.5142 | 5.0578 | 13.2971 | 0.1928 | 0.0205 | 0.0252 | 98.4035 | -7692 | 23001 | 49 |  |  | A99 |  | 12.9032 | 1 | 2019-08-13 06:18
2 / 1 . | 2.7256 | 11.9531 | 0.4488 | 9.0528 | 0.8444 | 4.0446 | 50.4531 | 4.7787 | 13.6856 | 0.1624 | 0.0291 | 0.0237 | 98.202 | -7659 | 23034 | 49 |  |  | ag2 |  | 12.9096 | 2 | 2019-08-13 06:26
3 / 1 . | 2.4276 | 11.8517 | 0.4289 | 9.1362 | 0.8538 | 4.0688 | 50.6389 | 4.9565 | 13.5503 | 0.1825 | 0.0229 | 0.0211 | 98.1394 | -7592 | 23007 | 49 |  |  | a99 |  | 12.8974 | 3 | 2019-08-13 06:35
recovery | 98.5643 | 98.0143 | 94.7253 | 101.0946 | 99.338 | 99.7974 | 99.3071 | 98.1489 | 100.9564 | 74.3707 | 80.5556 | 101.4493
Tuemidday
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.5896 | 11.9444 | 0.4493 | 9.0254 | 0.8523 | 4.0448 | 50.4659 | 4.9308 | 13.0296 | 0.1717 | 0.0292 | 0.0274 | 97.5604 | -7607 | 23004 | 49 |  |  | A99_Crystal3_run4 |  | 12.7672 | 1 | 2019-08-13 15:40
2 / 1 . | 2.6993 | 12.0782 | 0.4694 | 9.1483 | 0.8447 | 4.0998 | 50.7329 | 4.9008 | 13.4198 | 0.2245 | 0.0294 | 0.0198 | 98.6668 | -7589 | 23034 | 49 |  |  | A99_Crystal3_run5 |  | 12.9444 | 2 | 2019-08-13 15:48
3 / 1 . | 2.779 | 12.0355 | 0.4287 | 9.0258 | 0.9061 | 4.1007 | 51.0505 | 4.9243 | 13.3744 | 0.2238 | 0.029 | 0.022 | 98.8998 | -7666 | 23016 | 49 |  |  | A99_Crystal3_run6 |  | 12.9608 | 3 | 2019-08-13 15:57
recovery | 100.4595 | 98.8679 | 98.7106 | 100.382 | 100.7782 | 100.0188 | 99.7284 | 97.9027 | 99.19 | 85.7538 | 97.3333 | 100.2899
Wed night 1
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.8204 | 12.0988 | 0.4236 | 9.0617 | 0.8342 | 4.1388 | 50.014 | 4.9095 | 13.2058 | 0.2538 | 0.0262 | 0.0275 | 97.8142 | -7686 | 23004 | 34 |  |  | A99 |  | 12.8274 | 1 | 2019-08-14 20:09
2 / 1 . | 2.6157 | 11.7479 | 0.4386 | 9.0687 | 0.8527 | 4.1153 | 49.3745 | 4.8607 | 12.8139 | 0.188 | 0.0239 | 0.0233 | 96.1231 | -7656 | 23036 | 31 |  |  | ag2 |  | 12.596 | 2 | 2019-08-14 20:17
3 / 1 . | 2.6406 | 11.8316 | 0.4397 | 9.1523 | 0.8999 | 4.0614 | 49.8714 | 4.7615 | 13.5935 | 0.1569 | 0.0316 | 0.0211 | 97.4614 | -7587 | 22987 | 34 |  |  | a99 |  | 12.8295 | 3 | 2019-08-14 20:25
recovery | 100.569 | 97.8265 | 95.3773 | 100.689 | 100.1471 | 100.5922 | 97.7702 | 96.4152 | 98.6655 | 82.8077 | 90.7778 | 104.2029
Thur morning | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6369 | 11.6961 | 0.4403 | 9.0313 | 0.8387 | 4.0945 | 49.7317 | 4.7102 | 13.4534 | 0.2242 | 0.0272 | 0.0208 | 96.9054 | -7651 | 23000 | 34 |  |  | A99 |  | 12.7521 | 1 | 2019-08-15 08:10
2 / 1 . | 2.6933 | 11.748 | 0.4651 | 9.2309 | 0.8337 | 4.1107 | 49.4462 | 4.779 | 13.3166 | 0.1419 | 0.0302 | 0.0259 | 96.8215 | -7677 | 23036 | 33 |  |  | ag2 |  | 12.731 | 2 | 2019-08-15 08:18
3 / 1 . | 2.6768 | 11.8504 | 0.5257 | 9.0376 | 0.8307 | 4.0513 | 49.9177 | 4.9197 | 13.2453 | 0.2003 | 0.0322 | 0.0212 | 97.3088 | -7587 | 22982 | 37 |  |  | a99 |  | 12.7663 | 3 | 2019-08-15 08:26
recovery | 99.7012 | 96.7741 | 104.8425 | 100.7521 | 96.9067 | 100.1103 | 97.6626 | 95.6005 | 99.667 | 78.3402 | 99.5556 | 98.4058
Thur evening
DataSet/Point | Na2O | Al2O3 | P2O5 | CaO | K2O | TiO2 | SiO2 | MgO | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.7647 | 12.0529 | 0.389 | 9.0503 | 0.8828 | 4.0729 | 51.0278 | 4.7624 | 13.1558 | 0.2363 | 0.0265 | 0.0266 | 98.4481 | -7690 | 23007 | 40 |  |  | A99 |  | 12.8881 | 1 | 2019-08-15 18:12
2 / 1 . | 2.7469 | 12.0019 | 0.3858 | 9.1861 | 0.8858 | 4.052 | 50.204 | 4.9079 | 13.5796 | 0.1973 | 0.0277 | 0.0248 | 98.1999 | -7647 | 23038 | 42 |  |  | ag2 |  | 12.911 | 2 | 2019-08-15 18:20
3 / 1 . | 2.6691 | 11.9278 | 0.4378 | 9.1082 | 0.8755 | 4.1015 | 50.2603 | 4.8964 | 13.3359 | 0.2816 | 0.0267 | 0.0199 | 97.9407 | -7598 | 22981 | 39 |  |  | a99 |  | 12.8625 | 3 | 2019-08-15 18:29
recovery | 101.864 | 98.6609 | 88.8352 | 100.9175 | 102.3655 | 99.8644 | 99.2324 | 96.6474 | 99.8065 | 98.9212 | 89.8889 | 103.3333
Overall recovery | 99.7243 | 97.9845 | 97.9152 | 100.4901 | 100.8915 | 100.1103 | 98.68 | 97.135 | 99.3398 | 86.8247 | 91.2381 | 98.6957 | #REF!
Overall precision | 3.6784 | 0.90817 | 8.8485 | 0.87797 | 3.0197 | 0.6556 | 1.0213 | 1.6747 | 1.6531 | 17.7218 | 10.2632 | 12.7547
2020-01-01
Preferred | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl
| 2.677 | 12.157 | 0.455 | 50.888 | 5.024 | 0.861 | 9.032 | 4.081 | 13.383 | 0.241 | 0.03 | 0.023
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6068 | 12.0312 | 0.4672 | 50.0245 | 4.9468 | 0.8552 | 9.1199 | 4.0418 | 13.5027 | 0.1716 | 0.0292 | 0.0298 | 97.8267 | -6787 | 26763 | 77 |  |  | A99_Crystal3_1 |  | 12.8542 | 1 | 2020-01-28 13:57
av/Pref | 97.3777 | 98.9652 | 102.6813 | 98.3031 | 98.4634 | 99.3264 | 100.9732 | 99.0395 | 100.8944 | 71.2033 | 97.3333 | 129.5652 | 0 | 0
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.8039 | 11.7678 | 0.5388 | 49.3201 | 4.7912 | 0.9299 | 8.8702 | 3.9807 | 12.8508 | 0.198 | 0.0539 | 0.0417 | 96.147 | -6770 | 26739 | 77 |  |  | A99_crystal3_2 |  | 12.5906 | 1 | 2020-01-28 14:10
av/Pref | 104.7404 | 96.7986 | 118.4176 | 96.9189 | 95.3662 | 108.0023 | 98.2086 | 97.5423 | 96.0233 | 82.1577 | 179.6667 | 181.3043 | 0 | 0
2 / 1 . | 2.9626 | 12.0565 | 0.416 | 50.0285 | 4.9064 | 0.9236 | 8.9733 | 3.9926 | 13.4496 | 0.2429 | 0.0338 | 0.0218 | 98.0078 | -6768 | 26723 | 77 |  |  | A99_crystal3_3 |  | 12.8659 | 2 | 2020-01-28 14:19
av/Pref | 110.6687 | 99.1733 | 91.4286 | 98.311 | 97.6592 | 107.2706 | 99.3501 | 97.8339 | 100.4976 | 100.7884 | 112.6667 | 94.7826 | 0 | 0
4 / 1 . | [12 empty] | 0 | . | . | . |  |  | VG2_crystal2_2 |  | 0 | 4 | 2020-01-28 14:10
5 / 1 . | [12 empty] | 0 | . | . | . |  |  | VG2_crystal2_3 |  | 0 | 5 | 2020-01-28 14:10
Start after calibration
1 / 1 . | 2.6068 | 12.0312 | 0.4672 | 50.0245 | 4.9468 | 0.8552 | 9.1199 | 4.0418 | 13.5027 | 0.1716 | 0.0292 | 0.0298 | 97.8267 | -6787 | 26763 | 77 |  |  | A99_Crystal3_1 |  | 12.8542 | 1 | 2020-01-28 13:57
1 / 1 . | 2.8039 | 11.7678 | 0.5388 | 49.3201 | 4.7912 | 0.9299 | 8.8702 | 3.9807 | 12.8508 | 0.198 | 0.0539 | 0.0417 | 96.147 | -6770 | 26739 | 77 |  |  | A99_crystal3_2 |  | 12.5906 | 1 | 2020-01-28 14:10
2 / 1 . | 2.9626 | 12.0565 | 0.416 | 50.0285 | 4.9064 | 0.9236 | 8.9733 | 3.9926 | 13.4496 | 0.2429 | 0.0338 | 0.0218 | 98.0078 | -6768 | 26723 | 77 |  |  | A99_crystal3_3 |  | 12.8659 | 2 | 2020-01-28 14:19
av/Pref | 1.0426 | 0.98312 | 1.0418 | 0.97844 | 0.97163 | 1.0487 | 0.99511 | 0.98139 | 0.99138 | 0.84716 | 1.2989 | 1.3522 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 2020-01-28 13:00
Mon night standards | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.894 | 11.5591 | 0.4388 | 49.7297 | 4.9217 | 0.817 | 8.9181 | 4.092 | 12.9199 | 0.2184 | 0.0233 | 0.0241 | 96.556 | -6789 | 26731 | 77 |  |  | A99_crystal3_3 |  | 12.644 | 1 | 2020-01-29 02:15
2 / 1 . | 2.6417 | 11.5416 | 0.5287 | 49.3968 | 4.9387 | 0.85 | 8.8902 | 4.1077 | 13.3668 | 0.1643 | 0.0291 | 0.0268 | 96.4824 | -6765 | 26739 | 77 |  |  | A99_crystal3_2 |  | 12.6837 | 2 | 2020-01-29 02:23
3 / 1 hit crack | 2.682 | 11.1887 | 0.4165 | 48.9296 | 4.7505 | 0.8253 | 8.7478 | 4.0636 | 13.3938 | 0.1965 | 0.0497 | 0.0476 | 95.2916 | -6769 | 26700 | 77 |  |  | A99_crystal3_3 |  | 12.5515 | 3 | 2020-01-29 02:32
av/Pref | 1.0339 | 0.9501 | 1.0632 | 0.97397 | 0.98133 | 0.96806 | 0.98584 | 1.0046 | 0.98209 | 0.79398 | 0.87333 | 1.1065 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 2020-01-28 13:00
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.7911 | 12.441 | 0.5059 | 49.8513 | 4.8693 | 0.9836 | 8.8883 | 4.1174 | 12.9464 | 0.1974 | 0.0302 | 0.0224 | 97.6442 | -6620 | 26659 | 83 |  |  | A99_1 |  | 12.7722 | 1 | 2020-01-29 14:25
2 / 1 . | 2.8561 | 12.322 | 0.4351 | 50.0125 | 4.9284 | 0.8946 | 8.8953 | 4.0747 | 12.9545 | 0.2485 | 0.0317 | 0.0228 | 97.6763 | -6620 | 26644 | 83 |  |  | A99_2 |  | 12.7737 | 2 | 2020-01-29 14:33
3 / 1 . | [13 empty] | -6635 | 26624 | 83 |  |  | A99_3 |  | 0 | 3 | 2020-01-29 14:41
av/Pref | 1.0548 | 1.0185 | 1.0341 | 0.98121 | 0.97509 | 1.0907 | 0.98448 | 1.0037 | 0.96768 | 0.9251 | 1.0317 | 0.98261 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 2020-01-28 13:00
2020-02-01
| Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl
Preferred | 2.677 | 12.157 | 0.455 | 50.888 | 5.024 | 0.861 | 9.032 | 4.081 | 13.383 | 0.241 | 0.03 | 0.023
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6068 | 12.0312 | 0.4672 | 50.0245 | 4.9468 | 0.8552 | 9.1199 | 4.0418 | 13.5027 | 0.1716 | 0.0292 | 0.0298 | 97.8267 | -6787 | 26763 | 77 |  |  | A99_Crystal3_1 |  | 12.8542 | 1 | 43858.5813
av/Pref | 97.3777 | 98.9652 | 102.6813 | 98.3031 | 98.4634 | 99.3264 | 100.9732 | 99.0395 | 100.8944 | 71.2033 | 97.3333 | 129.5652 | 0 | 0
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.8039 | 11.7678 | 0.5388 | 49.3201 | 4.7912 | 0.9299 | 8.8702 | 3.9807 | 12.8508 | 0.198 | 0.0539 | 0.0417 | 96.147 | -6770 | 26739 | 77 |  |  | A99_crystal3_2 |  | 12.5906 | 1 | 43858.5905
av/Pref | 104.7404 | 96.7986 | 118.4176 | 96.9189 | 95.3662 | 108.0023 | 98.2086 | 97.5423 | 96.0233 | 82.1577 | 179.6667 | 181.3043 | 0 | 0
2 / 1 . | 2.9626 | 12.0565 | 0.416 | 50.0285 | 4.9064 | 0.9236 | 8.9733 | 3.9926 | 13.4496 | 0.2429 | 0.0338 | 0.0218 | 98.0078 | -6768 | 26723 | 77 |  |  | A99_crystal3_3 |  | 12.8659 | 2 | 43858.5964
av/Pref | 110.6687 | 99.1733 | 91.4286 | 98.311 | 97.6592 | 107.2706 | 99.3501 | 97.8339 | 100.4976 | 100.7884 | 112.6667 | 94.7826 | 0 | 0
4 / 1 . | [12 empty] | 0 | . | . | . |  |  |  |  | 0 | 4 | 43858.5901
5 / 1 . | [12 empty] | 0 | . | . | . |  |  |  |  | 0 | 5 | 43858.5903
Start after calibration
1 / 1 . | 2.6068 | 12.0312 | 0.4672 | 50.0245 | 4.9468 | 0.8552 | 9.1199 | 4.0418 | 13.5027 | 0.1716 | 0.0292 | 0.0298 | 97.8267 | -6787 | 26763 | 77 |  |  | A99_Crystal3_1 |  | 12.8542 | 1 | 43858.5813
1 / 1 . | 2.8039 | 11.7678 | 0.5388 | 49.3201 | 4.7912 | 0.9299 | 8.8702 | 3.9807 | 12.8508 | 0.198 | 0.0539 | 0.0417 | 96.147 | -6770 | 26739 | 77 |  |  | A99_crystal3_2 |  | 12.5906 | 1 | 43858.5905
2 / 1 . | 2.9626 | 12.0565 | 0.416 | 50.0285 | 4.9064 | 0.9236 | 8.9733 | 3.9926 | 13.4496 | 0.2429 | 0.0338 | 0.0218 | 98.0078 | -6768 | 26723 | 77 |  |  | A99_crystal3_3 |  | 12.8659 | 2 | 43858.5964
av/Pref | 1.0426 | 0.98312 | 1.0418 | 0.97844 | 0.97163 | 1.0487 | 0.99511 | 0.98139 | 0.99138 | 0.84716 | 1.2989 | 1.3522 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 43858.5418
Mon night standards | Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.894 | 11.5591 | 0.4388 | 49.7297 | 4.9217 | 0.817 | 8.9181 | 4.092 | 12.9199 | 0.2184 | 0.0233 | 0.0241 | 96.556 | -6789 | 26731 | 77 |  |  | A99_crystal3_3 |  | 12.644 | 1 | 43859.0937
2 / 1 . | 2.6417 | 11.5416 | 0.5287 | 49.3968 | 4.9387 | 0.85 | 8.8902 | 4.1077 | 13.3668 | 0.1643 | 0.0291 | 0.0268 | 96.4824 | -6765 | 26739 | 77 |  |  | A99_crystal3_2 |  | 12.6837 | 2 | 43859.0995
3 / 1 hit crack | 2.682 | 11.1887 | 0.4165 | 48.9296 | 4.7505 | 0.8253 | 8.7478 | 4.0636 | 13.3938 | 0.1965 | 0.0497 | 0.0476 | 95.2916 | -6769 | 26700 | 77 |  |  | A99_crystal3_3 |  | 12.5515 | 3 | 43859.1054
av/Pref | 1.0339 | 0.9501 | 1.0632 | 0.97397 | 0.98133 | 0.96806 | 0.98584 | 1.0046 | 0.98209 | 0.79398 | 0.87333 | 1.1065 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 43858.5418
Wed afternoon start
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.7911 | 12.441 | 0.5059 | 49.8513 | 4.8693 | 0.9836 | 8.8883 | 4.1174 | 12.9464 | 0.1974 | 0.0302 | 0.0224 | 97.6442 | -6620 | 26659 | 83 |  |  | A99_1 |  | 12.7722 | 1 | 43859.6004
2 / 1 . | 2.8561 | 12.322 | 0.4351 | 50.0125 | 4.9284 | 0.8946 | 8.8953 | 4.0747 | 12.9545 | 0.2485 | 0.0317 | 0.0228 | 97.6763 | -6620 | 26644 | 83 |  |  | A99_2 |  | 12.7737 | 2 | 43859.6062
3 / 1 . | 2.7972 | 12.2851 | 0.5081 | 49.4709 | 4.8449 | 0.8793 | 8.9495 | 4.0763 | 13.3809 | 0.1786 | 0.0239 | 0.0228 | 97.4175 | -6635 | 26624 | 83 |  |  | A99_3 |  | 12.7893 | 3 | 43859.6121
av/Pref | 1.0515 | 1.0158 | 1.0616 | 0.97819 | 0.97151 | 1.0676 | 0.98661 | 1.0021 | 0.9784 | 0.86376 | 0.95333 | 0.98551 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 43858.5418
Wed evening end
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6949 | 12.159 | 0.3706 | 50.4087 | 4.9403 | 0.8727 | 9.0593 | 4.0894 | 13.5223 | 0.2221 | 0.0236 | 0.0229 | 98.3857 | -6778 | 26755 | 83 |  |  | A99_crystal3_3 |  | 12.9209 | 1 | 43859.7588
2 / 1 . | 2.7456 | 12.022 | 0.5104 | 49.876 | 5.065 | 0.8722 | 9.055 | 4.0713 | 13.6511 | 0.1478 | 0.0217 | 0.0253 | 98.0633 | -6758 | 26720 | 82 |  |  | A99_crystal3_2 |  | 12.8914 | 2 | 43859.7647
3 / 1 . | 2.7687 | 12.1759 | 0.4485 | 50.3432 | 4.8456 | 0.8042 | 9.135 | 4.1184 | 13.4312 | 0.2191 | 0.0346 | 0.0245 | 98.3488 | -6816 | 26666 | 82 |  |  | A99_crystal3_3 |  | 12.9084 | 3 | 43859.7705
av/Pref | 1.0222 | 0.99687 | 0.97399 | 0.98666 | 0.98533 | 0.98688 | 1.0057 | 1.0029 | 1.0113 | 0.81466 | 0.88778 | 1.0536 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 43858.5418
Thur night start
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.7473 | 12.3886 | 0.3728 | 49.7576 | 4.9697 | 1.0483 | 8.9234 | 4.0603 | 13.3179 | 0.2114 | 0.028 | 0.0281 | 97.8535 | -6622 | 26668 | 83 |  |  | A99_1 |  | 12.8407 | 1 | 43860.7645
2 / 1 . | 2.9382 | 12.8942 | 0.4601 | 49.1648 | 4.967 | 0.8234 | 9.117 | 4.0488 | 12.8498 | 0.1774 | 0.0254 | 0.0221 | 97.4883 | -6612 | 26638 | 83 |  |  | A99_2 |  | 12.7395 | 2 | 43860.7703
3 / 1 . | 2.877 | 12.7329 | 0.4 | 49.6823 | 4.9116 | 0.8904 | 9.2023 | 4.0918 | 13.161 | 0.2351 | 0.0244 | 0.0262 | 98.235 | -6641 | 26621 | 83 |  |  | A99_3 |  | 12.8735 | 3 | 43860.7762
av/Pref | 1.0662 | 1.0424 | 0.90322 | 0.97341 | 0.98516 | 1.0693 | 1.0054 | 0.99656 | 0.97957 | 0.86293 | 0.86444 | 1.1072 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 43858.5418
Frinight
Oxide
DataSet/Point | Na2O | Al2O3 | P2O5 | SiO2 | MgO | K2O | CaO | TiO2 | FeO | MnO | SO2 | Cl | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 2.6749 | 11.9245 | 0.3942 | 49.1155 | 5.0117 | 0.8779 | 8.8359 | 4.0781 | 12.553 | 0.1711 | 0.0294 | 0.0264 | 95.6925 | -6765 | 26755 | 83 |  |  | A99_crystal3_3 |  | 12.5035 | 1 | 43861.774
2 / 1 . | 2.7177 | 11.9834 | 0.506 | 50.0468 | 4.9239 | 0.8934 | 8.9921 | 4.1082 | 13.1355 | 0.2641 | 0.0289 | 0.0272 | 97.6272 | -6756 | 26716 | 84 |  |  | A99_crystal3_2 |  | 12.7994 | 2 | 43861.7799
3 / 1 . | 2.7871 | 12.2975 | 0.418 | 50.1114 | 4.9217 | 0.8212 | 8.9044 | 4.0717 | 13.3719 | 0.1861 | 0.0326 | 0.0247 | 97.9483 | -6738 | 26630 | 87 |  |  | A99_crystal3_3 |  | 12.8396 | 3 | 43861.7857
av/Pref | 1.0185 | 0.99272 | 0.96571 | 0.97779 | 0.98576 | 1.0037 | 0.98658 | 1.0012 | 0.97289 | 0.85934 | 1.01 | 1.1348 | #DIV/0! | . | . | . |  |  |  |  | 0 | 4 | 43858.5418

## EPMA secondary standards - SCOl
2019-08-01
preferred | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total
| 48.809 | 41.218 | 9.795 | 0.037 | 0.108 | 0.149 | 0.363
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
4 / 1 . | 47.8167 | 40.5438 | 9.7621 | 0.035 | 0.138 | 0.1353 | 0.3975 | 98.8282 | -11069 | 31844 | 26 |  |  | SCOl_Crystal3_run1 |  | 11.6556 | 1 | 2019-08-12 13:02
5 / 1 . | 47.7136 | 40.6846 | 9.5333 | 0.0354 | 0.1415 | 0.1383 | 0.3535 | 98.6001 | -11064 | 31851 | 26 |  |  | SCOl_3_2 |  | 11.6006 | 2 | 2019-08-12 13:09
6 / 1 . | 47.7742 | 40.2498 | 9.624 | 0.0351 | 0.1311 | 0.1483 | 0.3869 | 98.3492 | -11057 | 31854 | 26 |  |  | SCOl_3_3 |  | 11.5882 | 3 | 2019-08-12 13:15
rec | 97.8675 | 98.2404 | 98.4155 | 95.045 | 126.7284 | 94.3848 | 104.4904 | #DIV/0!
pre laying test
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 47.8455 | 40.4687 | 9.8476 | 0.0314 | 0.1342 | 0.149 | 0.3828 | 98.8592 | -11083 | 31866 | 26 |  |  | SC_tueafter |  | 11.6677 | 1 | 2019-08-13 16:08
rec | 98.026 | 98.1821 | 100.537 | 84.8649 | 124.2593 | 100 | 105.4545 | #DIV/0!
sancarlostuenight1 | Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 47.9886 | 40.6964 | 9.6656 | 0.0341 | 0.1478 | 0.1465 | 0.3823 | 99.0613 | -11091 | 31866 | 29 |  |  | SC_tueafter |  | 11.6691 | 1 | 2019-08-13 17:41
2 / 1 . | 48.3228 | 40.7536 | 9.8418 | 0.0377 | 0.1404 | 0.1643 | 0.3686 | 99.6291 | -11108 | 31867 | 29 |  |  | SC_tueafter |  | 11.7485 | 2 | 2019-08-13 17:47
3 / 1 . | 48.4619 | 41.5479 | 9.7801 | 0.0332 | 0.134 | 0.1443 | 0.3756 | 100.4769 | -11116 | 31870 | 29 |  |  | SC_tueafter |  | 11.8311 | 3 | 2019-08-13 17:54
rec | 98.8706 | 99.4694 | 99.6682 | 94.5946 | 130.3086 | 101.8121 | 103.4435 | #DIV/0!
tuenight2 | Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 48.1419 | 40.5006 | 9.798 | 0.036 | 0.1266 | 0.1668 | 0.3802 | 99.1501 | -11052 | 31862 | 29 |  |  | SC_tueafter |  | 11.6935 | 1 | 2019-08-13 22:15
2 / 1 . | 48.336 | 40.9597 | 9.7696 | 0.0367 | 0.1355 | 0.1637 | 0.3531 | 99.7542 | -11059 | 31840 | 29 |  |  | SC_tueafter |  | 11.7515 | 2 | 2019-08-13 22:22
3 / 1 . | 48.543 | 41.2371 | 9.3974 | 0.0364 | 0.135 | 0.1486 | 0.3516 | 99.849 | -11116 | 31881 | 29 |  |  | SC_tueafter |  | 11.7176 | 3 | 2019-08-13 22:28
rec | 99.0397 | 99.2264 | 98.5707 | 98.2883 | 122.5617 | 107.1812 | 99.6235 | #DIV/0!
tuesnight3 | Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 47.7354 | 40.4252 | 9.7055 | 0.0362 | 0.1245 | 0.1498 | 0.3546 | 98.5313 | -11047 | 31862 | 29 |  |  | SC_tueafter |  | 11.6127 | 1 | 2019-08-14 01:04
2 / 1 . | 47.9571 | 40.9347 | 9.6754 | 0.0333 | 0.1267 | 0.1524 | 0.3585 | 99.2381 | -11047 | 31854 | 29 |  |  | SC_tueafter |  | 11.6858 | 2 | 2019-08-14 01:10
3 / 1 . | 48.3165 | 40.789 | 9.5841 | 0.0331 | 0.1233 | 0.1389 | 0.3741 | 99.359 | -11046 | 31847 | 29 |  |  | SC_tueafter |  | 11.6876 | 3 | 2019-08-14 01:17
rec | 98.3487 | 98.7828 | 98.5707 | 92.4324 | 115.5864 | 98.6801 | 99.8347 | #DIV/0!
tuenight4 | Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 48.834 | 41.3132 | 9.7114 | 0.0329 | 0.1339 | 0.1464 | 0.3571 | 100.5289 | -11096 | 31812 | 29 |  |  | SC_tueafter |  | 11.8254 | 1 | 2019-08-14 03:33
2 / 1 . | 49.0022 | 41.4286 | 9.5586 | 0.034 | 0.1311 | 0.1495 | 0.3525 | 100.6564 | -11096 | 31805 | 29 |  |  | SC_tueafter |  | 11.821 | 2 | 2019-08-14 03:40
3 / 1 . | 49.044 | 41.6149 | 9.6404 | 0.0355 | 0.1354 | 0.1484 | 0.369 | 100.9875 | -11096 | 31799 | 29 |  |  | SC_tueafter |  | 11.868 | 3 | 2019-08-14 03:46
rec | 100.3095 | 100.5683 | 98.3849 | 92.2523 | 123.5802 | 99.396 | 99.045 | #DIV/0!
wed morn 1st check | Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 48.1538 | 40.6035 | 9.6553 | 0.0357 | 0.1252 | 0.1559 | 0.3541 | 99.0836 | -11067 | 31789 | 27 |  |  | SC_tueafter |  | 11.6658 | 1 | 2019-08-14 08:00
2 / 1 . | 47.8959 | 40.7725 | 9.5447 | 0.0344 | 0.121 | 0.152 | 0.3535 | 98.874 | -11065 | 31783 | 27 |  |  | SC_tueafter |  | 11.631 | 2 | 2019-08-14 08:07
rec | 98.3934 | 98.7142 | 98.0092 | 94.7297 | 113.9815 | 103.3221 | 97.4656 | #DIV/0!
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 48.5633 | 41.4104 | 9.6059 | 0.0323 | 0.1209 | 0.1445 | 0.3745 | 100.2518 | -11079 | 31789 | 27 |  |  | SC_tueafter |  | 11.786 | 1 | 2019-08-14 11:41
2 / 1 . | 48.5865 | 40.8287 | 9.5824 | 0.0323 | 0.1215 | 0.1529 | 0.3559 | 99.6602 | -11067 | 31780 | 27 |  |  | SC_tueafter |  | 11.7179 | 2 | 2019-08-14 11:47
3 / 1 . | 49.0528 | 41.3875 | 9.4698 | 0.0359 | 0.1233 | 0.1737 | 0.3616 | 100.6046 | -11074 | 31775 | 27 |  |  | SC_tueafter |  | 11.8085 | 3 | 2019-08-14 11:54
rec | 99.8467 | 99.9778 | 97.5263 | 90.5405 | 112.8704 | 105.3915 | 100.2755 | #DIV/0!
overall recovery | 98.9363 | 99.2574 | 98.5697 | 93.5135 | 121.2919 | 101.2943 | 100.9773 | #DIV/0!
Overal precision | 0.89681 | 0.97568 | 1.2166 | 4.7532 | 5.6189 | 6.2084 | 3.6771
2020-01-01
preferred | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total
| 48.809 | 41.218 | 9.795 | 0.037 | 0.108 | 0.149 | 0.363
Mon afternoon | Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 47.9026 | 40.7672 | 9.5455 | 0.0432 | 0.1503 | 0.1307 | 0.3742 | 98.9138 | -14308 | 32518 | 34 |  |  | SC_Crystal3_1 |  | 11.6375 | 1 | 43858.6272
2 / 1 . | 47.9974 | 40.1448 | 9.6262 | 0.0348 | 0.1208 | 0.1308 | 0.3765 | 98.4312 | -14248 | 32560 | 34 |  |  | SC_Crystal3_2 |  | 11.5927 | 2 | 43858.6316
3 / 1 . | 47.9236 | 40.1144 | 9.8248 | 0.0392 | 0.1188 | 0.1118 | 0.373 | 98.5055 | -14237 | 32566 | 34 |  |  | SC_Crystal3_3 |  | 11.6206 | 3 | 43858.6361
rec | 98.222 | 97.875 | 98.6779 | 105.5856 | 120.3395 | 83.5123 | 103.1864
Monday evening after run
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 48.829 | 40.765 | 9.7486 | 0.0383 | 0.1429 | 0.1374 | 0.3364 | 99.9976 | -14300 | 32519 | 34 |  |  | SC_Crystal3_1 |  | 11.7691 | 1 | 43858.8747
2 / 1 . | 47.7383 | 39.9939 | 9.4544 | 0.0348 | 0.1129 | 0.1346 | 0.3427 | 97.8117 | -14239 | 32560 | 34 |  |  | SC_Crystal3_2 |  | 11.5031 | 2 | 43858.8792
3 / 1 . | 47.0595 | 39.6219 | 9.7536 | 0.0326 | 0.1096 | 0.1377 | 0.3482 | 97.0632 | -14228 | 32571 | 34 |  |  | SC_Crystal3_3 |  | 11.4592 | 3 | 43858.8837
rec | 98.0876 | 97.3529 | 98.5421 | 95.2252 | 112.7778 | 91.6555 | 94.3343
2020-02-01
preferred | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total
| 48.809 | 41.218 | 9.795 | 0.037 | 0.108 | 0.149 | 0.363
Mon afternoon | Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 47.9026 | 40.7672 | 9.5455 | 0.0432 | 0.1503 | 0.1307 | 0.3742 | 98.9138 | -14308 | 32518 | 34 |  |  | SC_Crystal3_1 |  | 11.6375 | 1 | 43858.6272
2 / 1 . | 47.9974 | 40.1448 | 9.6262 | 0.0348 | 0.1208 | 0.1308 | 0.3765 | 98.4312 | -14248 | 32560 | 34 |  |  | SC_Crystal3_2 |  | 11.5927 | 2 | 43858.6316
3 / 1 . | 47.9236 | 40.1144 | 9.8248 | 0.0392 | 0.1188 | 0.1118 | 0.373 | 98.5055 | -14237 | 32566 | 34 |  |  | SC_Crystal3_3 |  | 11.6206 | 3 | 43858.6361
rec | 98.222 | 97.875 | 98.6779 | 105.5856 | 120.3395 | 83.5123 | 103.1864
Monday evening after run
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 48.829 | 40.765 | 9.7486 | 0.0383 | 0.1429 | 0.1374 | 0.3364 | 99.9976 | -14300 | 32519 | 34 |  |  | SC_Crystal3_1 |  | 11.7691 | 1 | 43858.8747
2 / 1 . | 47.7383 | 39.9939 | 9.4544 | 0.0348 | 0.1129 | 0.1346 | 0.3427 | 97.8117 | -14239 | 32560 | 34 |  |  | SC_Crystal3_2 |  | 11.5031 | 2 | 43858.8792
3 / 1 . | 47.0595 | 39.6219 | 9.7536 | 0.0326 | 0.1096 | 0.1377 | 0.3482 | 97.0632 | -14228 | 32571 | 34 |  |  | SC_Crystal3_3 |  | 11.4592 | 3 | 43858.8837
rec | 98.0876 | 97.3529 | 98.5421 | 95.2252 | 112.7778 | 91.6555 | 94.3343
Tue evening run 1
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 48.3258 | 40.7178 | 9.883 | 0.0351 | 0.1092 | 0.1433 | 0.3428 | 99.5569 | -14298 | 32511 | 37 |  |  | SC_Crystal3_1 |  | 11.738 | 1 | 43859.794
2 / 1 . | 47.4272 | 40.019 | 9.5493 | 0.0301 | 0.103 | 0.1413 | 0.3673 | 97.6373 | -14239 | 32552 | 34 |  |  | SC_Crystal3_2 |  | 11.4994 | 2 | 43859.7986
3 / 1 . | 47.1363 | 39.2303 | 9.4358 | 0.0318 | 0.1078 | 0.1416 | 0.3544 | 96.438 | -14227 | 32562 | 34 |  |  | SC_Crystal3_3 |  | 11.3569 | 3 | 43859.8031
rec | 97.584 | 97.0184 | 98.2409 | 87.3874 | 98.7654 | 95.3468 | 97.7502
Tue evening run 2
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 49.0762 | 40.8319 | 9.5819 | 0.0341 | 0.1132 | 0.1142 | 0.3289 | 100.0802 | -14308 | 32505 | 37 |  |  | SC_Crystal3_1 |  | 11.7533 | 1 | 43860.0062
2 / 1 . | 46.8138 | 39.5447 | 9.6845 | 0.0319 | 0.1086 | 0.1227 | 0.3441 | 96.6503 | -14227 | 32562 | 34 |  |  | SC_Crystal3_2 |  | 11.4057 | 2 | 43860.0107
3 / 1 . | 48.3427 | 41.0027 | 9.7145 | 0.0335 | 0.127 | 0.1353 | 0.3336 | 99.6892 | -14238 | 32522 | 34 |  |  | SC_Crystal3_3 |  | 11.7324 | 3 | 43860.0153
rec | 98.5014 | 98.1604 | 98.6248 | 89.6396 | 107.6543 | 83.2662 | 92.4334
Tue evening run 3
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 48.8945 | 40.7851 | 9.8801 | 0.0318 | 0.1307 | 0.1395 | 0.324 | 100.1858 | -14327 | 32557 | 37 |  |  | SC_Crystal3_3 |  | 11.8018 | 1 | 43860.2588
2 / 1 . | 48.4411 | 40.8606 | 9.6526 | 0.0322 | 0.1389 | 0.1442 | 0.3404 | 99.6101 | -14325 | 32575 | 37 |  |  | SC_Crystal3_4 |  | 11.719 | 2 | 43860.2633
3 / 1 . | 48.3152 | 40.5392 | 9.5274 | 0.0347 | 0.1386 | 0.138 | 0.3311 | 99.0242 | -14329 | 32583 | 37 |  |  | SC_Crystal3_5 |  | 11.6403 | 3 | 43860.2678
rec | 99.4699 | 98.8119 | 98.8943 | 88.9189 | 125.9877 | 94.34 | 91.4141
Thur evening run
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 49.0927 | 40.785 | 9.4153 | 0.0338 | 0.108 | 0.1376 | 0.3309 | 99.9033 | -14291 | 32505 | 37 |  |  | SC_Crystal3_1 |  | 11.7179 | 1 | 43861.0718
2 / 1 . | 47.6591 | 39.634 | 9.5508 | 0.0337 | 0.1116 | 0.1389 | 0.3412 | 97.4694 | -14227 | 32530 | 34 |  |  | SC_Crystal3_2 |  | 11.4774 | 2 | 43861.0763
3 / 1 . | 48.4147 | 41.0154 | 9.6071 | 0.036 | 0.1262 | 0.1537 | 0.3447 | 99.6978 | -14243 | 32515 | 34 |  |  | SC_Crystal3_3 |  | 11.7243 | 3 | 43861.0808
| 99.1392 | 98.205 | 97.2374 | 93.2432 | 106.7284 | 96.2416 | 93.3701
Sat afternoon run
Oxide
DataSet/Point | MgO | SiO2 | FeO | Al2O3 | CaO | MnO | NiO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 47.8621 | 40.1828 | 9.6392 | 0.032 | 0.1219 | 0.1324 | 0.3263 | 98.2967 | -14259 | 32505 | 40 |  |  | SC_Crystal3_1 |  | 11.5739 | 1 | 43862.7409
2 / 1 . | 47.9215 | 40.2193 | 9.7625 | 0.0353 | 0.1488 | 0.1358 | 0.3413 | 98.5644 | -14271 | 32530 | 41 |  |  | SC_Crystal3_2 |  | 11.6202 | 2 | 43862.7455
rec | 98.1208 | 97.5328 | 99.0388 | 90.9459 | 125.3241 | 90 | 91.9559
Overall std dev | 0.63447 | 0.52816 | 0.13551 | 0.0029476 | 0.014435 | 0.0095304 | 0.015701
Mean | 48.0587 | 40.3388 | 9.6419 | 0.034445 | 0.12244 | 0.13508 | 0.3451
Accuracy | 98.4627 | 97.8668 | 98.4365 | 93.0946 | 113.3704 | 90.6544 | 95.0689
Precision | 1.3202 | 1.3093 | 1.4054 | 8.5575 | 11.7896 | 7.0556 | 4.5498

## EPMA secondary standards _ Aug8
2019-08-01
| Na2O | Al2O3 | TiO2 | CaO | SiO2 | MgO | FeO | Cr2O3 | MnO
Prefered | 1.2848 | 8.379 | 0.88136 | 16.1813 | 50.3705 | 15.9974 | 6.422 | N.d. | 0.158
Oxide
DataSet/Point | Na2O | Al2O3 | TiO2 | CaO | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 1.2841 | 8.2846 | 0.8684 | 15.9171 | 50.6338 | 15.7476 | 6.3523 | 0.1691 | 0.139 | 99.3959 | 8742 | 32579 | -28 |  |  | Hole8_aug_crystal1_1 |  | 12.363 | 1 | 2019-08-14 12:10
rec | 99.9437 | 98.8732 | 98.5294 | 98.3674 | 100.5227 | 98.4383 | 98.9141 | #VALUE! | 87.9747 | #DIV/0! | #DIV/0! | #DIV/0!
before px run | Oxide
DataSet/Point | Na2O | Al2O3 | TiO2 | CaO | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 1.2829 | 8.3167 | 0.8557 | 15.6848 | 50.3238 | 15.8784 | 6.3859 | 0.1559 | 0.1006 | 98.9848 | 8734 | 32579 | -28 |  |  | Hole8_aug_crystal1_1 |  | 12.3026 | 1 | 2019-08-14 12:46
2 / 1 . | 1.3344 | 8.2955 | 0.852 | 15.6766 | 50.5582 | 15.6677 | 6.4343 | 0.1516 | 0.1709 | 99.1413 | 8742 | 32579 | -28 |  |  | Hole8_aug_crystal1_2 |  | 12.3317 | 2 | 2019-08-14 12:52
3 / 1 . | 1.3761 | 8.3207 | 0.8603 | 15.8343 | 50.4244 | 15.865 | 6.4089 | 0.1525 | 0.1248 | 99.3669 | 8742 | 32584 | -28 |  |  | Hole8_aug_crystal1_3 |  | 12.3571 | 3 | 2019-08-14 12:58
rec | 103.6044 | 99.1879 | 97.1225 | 97.2229 | 100.1289 | 98.789 | 99.8079 | #VALUE! | 83.6076 | #DIV/0! | #DIV/0! | #DIV/0!
after px run
Oxide
DataSet/Point | Na2O | Al2O3 | TiO2 | CaO | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 1.2903 | 8.3223 | 0.8591 | 15.833 | 50.6791 | 15.8958 | 6.251 | 0.1755 | 0.1634 | 99.4694 | 8725 | 32579 | -28 |  |  | Hole8_aug_crystal1_1 |  | 12.3566 | 1 | 2019-08-14 17:17
2 / 1 . | 1.3584 | 8.2628 | 0.8549 | 15.9294 | 50.3245 | 15.8963 | 6.4802 | 0.1373 | 0.1206 | 99.3644 | 8742 | 32570 | -28 |  |  | Hole8_aug_crystal1_2 |  | 12.3684 | 2 | 2019-08-14 17:23
3 / 1 . | 1.3237 | 8.3561 | 0.8562 | 15.9787 | 50.8861 | 15.8573 | 6.1584 | 0.1574 | 0.184 | 99.7579 | 8735 | 32594 | -28 |  |  | Hole8_aug_crystal1_3 |  | 12.3862 | 3 | 2019-08-14 17:29
rec | 103.0596 | 99.2209 | 97.2057 | 98.3464 | 100.5149 | 99.2855 | 98.0457 | #VALUE! | 98.7342 | #DIV/0! | #DIV/0! | #DIV/0!
Overall recovery | 102.848 | 99.1571 | 97.3591 | 97.8679 | 100.3506 | 98.9517 | 98.925 | #VALUE! | 90.7143
overall precision | 2.6178 | 0.3374 | 0.57219 | 0.69151 | 0.37719 | 0.51304 | 1.6322 | 7.3298 | 19.5952
2020-02-01
| Na2O | Al2O3 | CaO | TiO2 | SiO2 | MgO | FeO | Cr2O3 | MnO
| 1.285 | 8.379 | 16.181 | 0.881 | 50.371 | 15.997 | 6.422 |  | 0.158
Oxide
DataSet/Point | Na2O | Al2O3 | CaO | TiO2 | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 1.352 | 8.5094 | 15.9568 | 0.887 | 49.6698 | 15.7084 | 6.2928 | 0.1484 | 0.1325 | 98.6573 | 8152 | 32318 | -50 |  |  | Augite_hol8_Crysatl1 |  | 12.2769 | 1 | 43860.7304
rec | 105.214 | 101.5563 | 98.6144 | 100.681 | 98.6079 | 98.1959 | 97.9882 | #DIV/0! | 83.8608 | #DIV/0!
Oxide
DataSet/Point | Na2O | Al2O3 | CaO | TiO2 | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 1.3247 | 8.0876 | 15.9227 | 0.8689 | 49.8274 | 15.8898 | 6.2961 | 0.1693 | 0.1359 | 98.5225 | 8145 | 32301 | -50 |  |  | Augite_hol8_Crysatl1 |  | 12.2619 | 1 | 43861.0853
2 / 1 . | 1.3792 | 8.1863 | 15.8625 | 0.8593 | 50.0046 | 15.8598 | 6.3371 | 0.1573 | 0.1251 | 98.7712 | 8125 | 32288 | -50 |  |  | Augite_hol8_Crysatl2 |  | 12.2869 | 2 | 43861.0896
3 / 1 . | 1.4069 | 8.3725 | 15.8669 | 0.8566 | 49.7625 | 15.8603 | 6.407 | 0.1557 | 0.1314 | 98.8198 | 8153 | 32288 | -50 |  |  | Augite_hol8_Crysatl3 |  | 12.3001 | 3 | 43861.0939
rec | 106.6355 | 98.0483 | 98.1647 | 97.798 | 98.9951 | 99.2059 | 98.828 | #DIV/0! | 82.7848 | #DIV/0!
Sat Afternoon run
Oxide
DataSet/Point | Na2O | Al2O3 | CaO | TiO2 | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 1.358 | 8.1885 | 15.9359 | 0.8752 | 49.9221 | 15.8461 | 6.3971 | 0.178 | 0.133 | 98.8339 | 8063 | 32284 | -43 |  |  | Augite_hol8_Crysatl1_1 |  | 12.3082 | 1 | 43862.7324
2 / 1 . | 1.3238 | 8.2013 | 15.909 | 0.8749 | 49.792 | 15.8382 | 6.2652 | 0.1498 | 0.1658 | 98.5199 | 8034 | 32326 | -44 |  |  | Augite_hol8_Crysatl1_2 |  | 12.2592 | 2 | 43862.7366
rec | 104.3502 | 97.8028 | 98.4021 | 99.3246 | 98.9797 | 99.032 | 98.5853 | #DIV/0! | 94.557 | #DIV/0!
Sun Afternoon
Oxide
DataSet/Point | Na2O | Al2O3 | CaO | TiO2 | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 1.3459 | 8.1507 | 15.9618 | 0.8674 | 49.2859 | 15.5532 | 6.4317 | 0.1335 | 0.1636 | 97.8936 | 7982 | 32326 | -50 |  |  | Augite_hol8_Crysatl1_1 |  | 12.2123 | 1 | 43863.6706
2 / 1 . | 1.3286 | 8.249 | 15.8092 | 0.8675 | 49.6709 | 15.6789 | 6.3304 | 0.1514 | 0.1402 | 98.2261 | 8044 | 32242 | -50 |  |  | Augite_hol8_Crysatl1_2 |  | 12.2266 | 2 | 43863.6749
3 / 1 . | 1.3101 | 8.1629 | 15.8207 | 0.879 | 49.5034 | 15.7399 | 6.3123 | 0.1362 | 0.1019 | 97.9663 | 7989 | 32324 | -50 |  |  | Augite_hol8_Crysatl1_3 |  | 12.1926 | 3 | 43863.6792
rec | 103.3619 | 97.7149 | 98.0403 | 98.899 | 98.2445 | 97.8767 | 99.0055 | #DIV/0! | 85.5907 | #DIV/0!
Oxide
DataSet/Point | Na2O | Al2O3 | CaO | TiO2 | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 1.3766 | 8.1365 | 15.8353 | 0.864 | 49.8732 | 15.7772 | 6.2595 | 0.1555 | 0.0903 | 98.3683 | 8010 | 32242 | -50 |  |  | Augite_hol8_Crysatl1_2 |  | 12.23 | 1 | 43863.9092
2 / 1 . | 1.3597 | 8.1584 | 15.8925 | 0.8689 | 49.4185 | 15.9693 | 6.3845 | 0.1512 | 0.1133 | 98.3164 | 8021 | 32240 | -50 |  |  | Augite_hol8_Crysatl1_3 |  | 12.2433 | 2 | 43863.9134
rec | 106.4708 | 97.2365 | 98.0403 | 98.3485 | 98.5604 | 99.2264 | 98.4429 | #DIV/0! | 64.4304 | #DIV/0!

## EPMA secondary standards _ Aug3
2019-08-01
DataSet/Point | Na2O | Al2O3 | TiO2 | CaO | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
| 0.82977 | 7.2704 | 0.48361 | 17.3614 | 50.979 | 17.2393 | 4.7888 | nd | 0.13852
Oxide
DataSet/Point | Na2O | Al2O3 | TiO2 | CaO | SiO2 | MgO | FeO | Cr2O3 | MnO | Total | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 0.8196 | 7.1892 | 0.48 | 16.973 | 50.3659 | 16.7699 | 4.8267 | 0.9211 | 0.1688 | 98.5144 | 19634 | 25937 | -34 |  |  | Augite_hole3_1 |  | 12.201 | 1 | 2019-08-14 13:05
2 / 1 . | 0.8756 | 7.1997 | 0.4696 | 17.0415 | 50.79 | 17.0065 | 4.7307 | 0.8486 | 0.1091 | 99.0713 | 19608 | 25937 | -34 |  |  | Augite_hole3_2 |  | 12.2405 | 2 | 2019-08-14 13:11
3 / 1 . | 0.846 | 7.2426 | 0.4635 | 16.9993 | 50.4809 | 16.852 | 4.8302 | 0.8934 | 0.1618 | 98.7697 | 19598 | 25937 | -34 |  |  | Augite_hole3_3 |  | 12.2261 | 3 | 2019-08-14 13:17
Overall recovery | 102.0849 | 99.176 | 97.3998 | 97.9447 | 99.1498 | 97.8931 | 100.1467 | #VALUE! | 105.8116
overall precision | 2.7004 | 0.32036 | 1.4462 | 0.16593 | 0.3543 | 0.58122 | 0.96128 | 3.365 | 18.1805

## EPMA secondary standards_ Labra
2020-02-01
| Na2O | SiO2 | MgO | Al2O3 | K2O | CaO | TiO2 | FeO | MnO | Total
| 3.716 | 51.573 | 0.128 | 30.167 | 0.131 | 13.566 | 0.042 | 0.436 | 0.012
Oxide
DataSet/Point | Na2O | SiO2 | MgO | Al2O3 | K2O | CaO | TiO2 | FeO | MnO | Total | X | Y | Z | An | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 3.8188 | 51.1788 | 0.1237 | 31.6047 | 0.1143 | 13.4477 | 0.0405 | 0.5298 | 0.025 | 100.8834 | 7342 | 22323 | -77 | 0.65617 |  |  | Lab_hole6_xstal1 |  | 11.6749 | 1 | 2020-01-31 09:04
2 / 1 . | 3.9466 | 50.691 | 0.1219 | 31.4588 | 0.1113 | 13.4438 | 0.0476 | 0.4349 | 0.0377 | 100.2936 | 7349 | 22323 | -77 | 0.6489 |  |  | Lab_hole6_xstal1_2 |  | 11.6013 | 2 | 2020-01-31 09:18
| 104.486 | 98.7627 | 95.9375 | 104.524 | 86.1069 | 99.1136 | 104.881 | 110.6307 | 261.25 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #VALUE!
Detection limit
Det.Lim ppm
DataSet/Point | Na | Si | Mg | Al | K | Ca | Ti | Fe | Mn | O | X | Y | Z |  | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 682 | 796 | 178 | 671 | 353 | 964 | 122 | 654 | 663 |  | 7342 | 22323 | -77 |  |  |  | Lab_hole6_xstal1 |  | 11.6749 | 1 | 2020-01-31 09:04
Actual ppm
| 37160 | 515730 | 1280 | 301670 | 1310 | 135660 | 420 | 4360 | 120 | 0 | 0
Actual run conditions
Oxide
DataSet/Point | Na2O | SiO2 | MgO | Al2O3 | K2O | CaO | TiO2 | FeO | MnO | Total | X | Y | Z |  | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 3.8554 | 51.04 | 0.1178 | 31.5293 | 0.1191 | 13.4674 | 0.0415 | 0.4417 | 0.0142 | 100.6264 | 7328 | 22323 | -77 | 0.65421 |  |  | Lab_hole6_xstal1 |  | 11.6376 | 1 | 2020-01-31 09:56
2 / 1 . | 3.9802 | 51.0317 | 0.1292 | 31.7111 | 0.1193 | 13.6072 | 0.0415 | 0.3914 | 0.0432 | 101.0548 | 7365 | 22322 | -77 | 0.64946 |  |  | Lab_hole6_xstal1_2 |  | 11.6883 | 2 | 2020-01-31 10:01
3 / 1 . | 3.8605 | 51.4519 | 0.1339 | 31.5712 | 0.1388 | 13.6088 | 0.0471 | 0.4323 | 0.0104 | 101.2548 | 7377 | 22331 | -77 | 0.65553 |  |  | Lab_hole6_xstal2 |  | 11.7136 | 3 | 2020-01-31 10:05
| 105.4992 | 99.3578 | 102.7734 | 104.8866 | 98.5115 | 100.3096 | 105.4762 | 94.461 | 223.3333 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #VALUE!
Oxide
DataSet/Point | Na2O | SiO2 | MgO | Al2O3 | K2O | CaO | TiO2 | FeO | MnO | Total | X | Y | Z |  | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 3.9355 | 50.6956 | 0.1254 | 29.6367 | 0.1321 | 13.829 | 0.0357 | 0.4447 | -0.018 | 98.8168 | 7390 | 22323 | -77 | 0.65516 |  |  | Lab_hole6_xstal1 |  | 11.4629 | 1 | 2020-02-01 01:54
2 / 1 . | 3.8853 | 50.9001 | 0.1176 | 30.7466 | 0.1494 | 13.802 | 0.0461 | 0.4678 | -0.018 | 100.0968 | 7353 | 22288 | -73 | 0.65691 |  |  | Lab_hole6_xstal1_2 |  | 11.6025 | 2 | 2020-02-01 01:58
3 / 1 . | 3.9414 | 50.9664 | 0.1287 | 30.0585 | 0.1099 | 13.6446 | 0.0409 | 0.4579 | 0.0119 | 99.3602 | 7364 | 22263 | -77 | 0.65261 |  |  | Lab_hole6_xstal2 |  | 11.5137 | 3 | 2020-02-01 02:03
| 105.5095 | 98.6059 | 96.7969 | 99.9346 | 99.5929 | 101.4192 | 97.381 | 104.7706 | -66.9444 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #VALUE!
Oxide
DataSet/Point | Na2O | SiO2 | MgO | Al2O3 | K2O | CaO | TiO2 | FeO | MnO | Total | X | Y | Z |  | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 4.0406 | 51.0474 | 0.1477 | 30.0396 | 0.1233 | 13.6 | 0.0414 | 0.479 | 0.0252 | 99.5443 | 7381 | 22323 | -74 | 0.64582 |  |  | Lab_hole6_xstal1 |  | 11.5351 | 1 | 2020-02-01 07:39
2 / 1 . | 3.6743 | 51.1715 | 0.1187 | 30.5728 | 0.1756 | 13.4401 | 0.0444 | 0.4457 | 0.0135 | 99.6565 | 7345 | 22288 | -72 | 0.66214 |  |  | Lab_hole6_xstal1_2 |  | 11.5379 | 2 | 2020-02-01 07:44
3 / 1 . | 3.9141 | 51.1694 | 0.1292 | 30.0165 | 0.1225 | 13.5104 | 0.0347 | 0.4028 | 0.0385 | 99.3381 | 7353 | 22257 | -77 | 0.65145 |  |  | Lab_hole6_xstal2 |  | 11.5009 | 3 | 2020-02-01 07:48
| 104.3147 | 99.1399 | 103.0208 | 100.1413 | 107.2265 | 99.6376 | 95.6349 | 101.4908 | 214.4444 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #VALUE!
Sat Night 1 | Det.Lim ppm Na: Oxide
DataSet/Point | Na2O | SiO2 | MgO | Al2O3 | K2O | CaO | TiO2 | FeO | MnO | Total | X | Y | Z |  | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 3.8005 | 50.387 | 0.1259 | 30.3704 | 0.1137 | 13.3361 | 0.0375 | 0.3943 | 0.0007 | 98.5661 | 7394 | 22323 | -74 | 0.65538 |  |  | Lab_hole6_xstal1 |  | 11.4022 | 1 | 2020-02-01 23:00
2 / 1 . | 3.7816 | 50.7946 | 0.1404 | 30.2318 | 0.1457 | 13.7854 | 0.0339 | 0.4861 | 0.0063 | 99.4058 | 7335 | 22284 | -74 | 0.6627 |  |  | Lab_hole6_xstal1_2 |  | 11.5319 | 2 | 2020-02-01 23:04
3 / 1 . | 4.062 | 50.6737 | 0.1192 | 30.1746 | 0.1564 | 13.1798 | 0.0317 | 0.4009 | -0.0072 | 98.7911 | 7366 | 22256 | -77 | 0.6362 |  |  | Lab_hole6_xstal2 |  | 11.4187 | 3 | 2020-02-01 23:09
| 104.4501 | 98.1491 | 100.3906 | 100.3047 | 105.8015 | 99.0253 | 81.8254 | 97.9587 | -0.55556 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #VALUE!
Sat Night 2 | Det.Lim ppm Na: Oxide
DataSet/Point | Na2O | SiO2 | MgO | Al2O3 | K2O | CaO | TiO2 | FeO | MnO | Total | X | Y | Z |  | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 3.8309 | 51.3717 | 0.1044 | 30.889 | 0.1334 | 13.5453 | 0.0438 | 0.4835 | 0.0248 | 100.4266 | 7375 | 22314 | -74 | 0.65638 |  |  | Lab_hole6_xstal1 |  | 11.6285 | 1 | 2020-02-02 03:27
2 / 1 . | 3.8923 | 50.834 | 0.1245 | 29.956 | 0.1062 | 13.3142 | 0.0455 | 0.4617 | -0.0421 | 98.6922 | 7335 | 22271 | -76 | 0.64998 |  |  | Lab_hole6_xstal1_2 |  | 11.4178 | 2 | 2020-02-02 03:31
3 / 1 . | 3.9831 | 51.3814 | 0.1467 | 30.5195 | 0.1045 | 13.3266 | 0.0478 | 0.4645 | 0.0023 | 99.9764 | 7357 | 22260 | -77 | 0.64508 |  |  | Lab_hole6_xstal2 |  | 11.5607 | 3 | 2020-02-02 03:36
| 105.0081 | 99.2684 | 97.8125 | 100.9541 | 87.5573 | 98.7422 | 108.8095 | 107.7752 | -41.6667 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #VALUE!
Sun Night
Oxide
DataSet/Point | Na2O | SiO2 | MgO | Al2O3 | K2O | CaO | TiO2 | FeO | MnO | Total | X | Y | Z |  | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
3 / 1 . | 3.8596 | 50.3495 | 0.1263 | 28.4981 | 0.1148 | 13.4906 | 0.0451 | 0.4674 | 0.0338 | 96.9852 | 7384 | 22267 | -84 | 0.65452 |  |  | Augite_hol8_Crysatl1_1 |  | 11.2551 | 1 | 2020-02-02 22:01
4 / 1 . | 3.8926 | 50.3153 | 0.1145 | 29.2582 | 0.1606 | 13.4902 | 0.046 | 0.4106 | -0.0133 | 97.6747 | 7384 | 22254 | -81 | 0.6509 |  |  | Augite_hol8_Crysatl1_2 |  | 11.3199 | 2 | 2020-02-02 22:06
5 / 1 . | 3.9889 | 50.6252 | 0.1369 | 29.1851 | 0.1335 | 13.2647 | 0.0475 | 0.4384 | -0.0151 | 97.8052 | 7370 | 22234 | -84 | 0.64261 |  |  | Augite_hol8_Crysatl1_3 |  | 11.3218 | 3 | 2020-02-02 22:11
| 105.3202 | 97.7837 | 98.3594 | 96.0668 | 104.0458 | 98.8882 | 110 | 100.6422 | 15 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  | #VALUE!

## Detection_limits_Glass
MountA part 2
Det.Lim ppm
DataSet/Point | Na | Al | P | Ca | K | Ti | Si | Mg | Fe | Mn | S | Cl | O | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 680 | 262 | 353 | 450 | 425 | 176 | 822 | 369 | 1345 | 1022 | 62 | 64 |  | 4913 | -2094 | 5 |  |  | A_LL3_G5_1 |  | 12.7933 | 1 | 2019-08-15 02:42
2 / 1 . | 668 | 250 | 361 | 449 | 426 | 177 | 813 | 387 | 1275 | 954 | 62 | 64 |  | 4894 | -2083 | 5 |  |  | A_LL3_G5_2 |  | 12.8533 | 2 | 2019-08-15 02:51
3 / 1 . | 602 | 255 | 324 | 439 | 432 | 178 | 850 | 384 | 1452 | 951 | 63 | 66 |  | 5802 | -896 | 7 |  |  | A_LL3_G6_1 |  | 12.7277 | 3 | 2019-08-15 02:59
4 / 1 . | 654 | 258 | 349 | 457 | 426 | 178 | 798 | 416 | 1355 | 940 | 62 | 65 |  | 5791 | -896 | 7 |  |  | A_LL3_G6_2 |  | 12.8645 | 4 | 2019-08-15 03:07
5 / 1 . | 667 | 257 | 334 | 442 | 422 | 176 | 815 | 384 | 1317 | 968 | 60 | 65 |  | 11483 | -1108 | -8 |  |  | A_LL3_G7_1 |  | 12.8279 | 5 | 2019-08-15 03:16
6 / 1 . | 674 | 256 | 307 | 433 | 398 | 176 | 788 | 412 | 1337 | 834 | 62 | 65 |  | 11483 | -1128 | -8 |  |  | A_LL3_G7_2 |  | 12.8485 | 6 | 2019-08-15 03:24
7 / 1 . | 623 | 253 | 346 | 430 | 423 | 172 | 798 | 385 | 1268 | 902 | 60 | 65 |  | 18314 | -1077 | -31 |  |  | A_LL3_G8_1 |  | 12.7763 | 7 | 2019-08-15 03:33
8 / 1 . | 596 | 254 | 360 | 431 | 451 | 174 | 808 | 406 | 1358 | 918 | 60 | 66 |  | 18307 | -1089 | -31 |  |  | A_LL3_G8_2 |  | 12.6987 | 8 | 2019-08-15 03:41
9 / 1 . | 696 | 262 | 340 | 419 | 420 | 172 | 823 | 397 | 1227 | 845 | 62 | 66 |  | 23420 | -1148 | -51 |  |  | A_LL3_G8b_1 |  | 12.7208 | 9 | 2019-08-15 03:49
10 / 1 . | 674 | 259 | 324 | 449 | 384 | 174 | 834 | 402 | 1176 | 889 | 62 | 64 |  | 23442 | -1148 | -51 |  |  | A_LL3_G8b_2 |  | 12.7972 | 10 | 2019-08-15 03:58
11 / 1 . | 596 | 255 | 346 | 454 | 411 | 172 | 799 | 406 | 1332 | 946 | 63 | 63 |  | 20511 | -3155 | -38 |  |  | A_LL3_G13_1 |  | 12.7549 | 11 | 2019-08-15 04:06
12 / 1 . | 651 | 255 | 364 | 454 | 403 | 176 | 822 | 398 | 1204 | 893 | 61 | 66 |  | 20491 | -3155 | -38 |  |  | A_LL3_G13_2 |  | 12.9365 | 12 | 2019-08-15 04:15
13 / 1 . | 663 | 259 | 341 | 437 | 412 | 174 | 807 | 408 | 1203 | 1025 | 62 | 65 |  | 19573 | -2939 | -35 |  |  | A_LL3_G12_1 |  | 12.8005 | 13 | 2019-08-15 04:23
14 / 1 . | 647 | 260 | 347 | 441 | 417 | 177 | 832 | 418 | 1179 | 1011 | 63 | 65 |  | 19568 | -2926 | -35 |  |  | A_LL3_G12_2 |  | 12.9134 | 14 | 2019-08-15 04:31
15 / 1 . | 623 | 255 | 372 | 455 | 394 | 173 | 842 | 379 | 1300 | 1108 | 61 | 67 |  | 9953 | -3507 | -6 |  |  | A_LL3_G9_1 |  | 12.7286 | 15 | 2019-08-15 04:40
16 / 1 . | 638 | 258 | 369 | 450 | 428 | 179 | 834 | 396 | 1311 | 1024 | 62 | 64 |  | 9938 | -3495 | -6 |  |  | A_LL3_G9_2 |  | 12.8449 | 16 | 2019-08-15 04:48
17 / 1 . | 593 | 255 | 336 | 444 | 391 | 165 | 816 | 432 | 1189 | 957 | 61 | 65 |  | 5167 | -7026 | -2 |  |  | A_LL4_G4_1 |  | 12.5998 | 17 | 2019-08-15 04:57
18 / 1 . | 588 | 259 | 372 | 457 | 389 | 164 | 876 | 404 | 1322 | 1078 | 62 | 66 |  | 5167 | -7015 | -2 |  |  | A_LL4_G4_2 |  | 12.5305 | 18 | 2019-08-15 05:05
19 / 1 . | 585 | 258 | 317 | 461 | 381 | 166 | 853 | 432 | 1268 | 774 | 62 | 65 |  | 8674 | -7252 | -2 |  |  | A_LL4_G5_1 |  | 12.6751 | 19 | 2019-08-15 05:13
20 / 1 . | 615 | 268 | 319 | 450 | 384 | 166 | 793 | 404 | 1254 | 1064 | 64 | 65 |  | 12208 | -7252 | -11 |  |  | A_LL4_G6_1 |  | 12.6257 | 20 | 2019-08-15 05:22
21 / 1 . | 602 | 257 | 327 | 466 | 402 | 166 | 816 | 418 | 1206 | 821 | 62 | 65 |  | 12187 | -7221 | -11 |  |  | A_LL4_G6_2 |  | 12.7182 | 21 | 2019-08-15 05:30
22 / 1 . | 587 | 257 | 299 | 458 | 374 | 165 | 795 | 416 | 1250 | 894 | 59 | 64 |  | 15860 | -6188 | -25 |  |  | A_LL4_G7_1 |  | 12.6946 | 22 | 2019-08-15 05:39
23 / 1 . | 569 | 264 | 327 | 452 | 388 | 166 | 825 | 412 | 1386 | 975 | 62 | 64 |  | 15879 | -6176 | -25 |  |  | A_LL4_G7_2 |  | 12.6602 | 23 | 2019-08-15 05:47
24 / 1 . | 605 | 265 | 334 | 457 | 357 | 167 | 812 | 407 | 1201 | 950 | 62 | 64 |  | 19669 | -6402 | -40 |  |  | A_LL4_G8_1 |  | 12.5293 | 24 | 2019-08-15 05:56
25 / 1 . | 623 | 258 | 388 | 445 | 342 | 166 | 835 | 422 | 1226 | 1017 | 62 | 66 |  | 19669 | -6417 | -40 |  |  | A_LL4_G8_2 |  | 12.4828 | 25 | 2019-08-15 06:04
26 / 1 . | 625 | 266 | 307 | 473 | 378 | 171 | 860 | 398 | 1303 | 901 | 61 | 65 |  | 19172 | -9138 | -40 |  |  | A_LL4_G12_1 |  | 12.5292 | 26 | 2019-08-15 06:12
27 / 1 . | 645 | 261 | 360 | 450 | 396 | 168 | 845 | 417 | 1303 | 1009 | 63 | 66 |  | 19172 | -9087 | -40 |  |  | A_LL4_G12_2 |  | 12.4315 | 27 | 2019-08-15 06:21
28 / 1 . | 634 | 263 | 352 | 466 | 395 | 168 | 824 | 423 | 1251 | 904 | 62 | 65 |  | 10595 | -9603 | -10 |  |  | A_LL4_G10_1 |  | 12.6649 | 28 | 2019-08-15 06:29
29 / 1 . | 540 | 269 | 363 | 483 | 408 | 167 | 845 | 425 | 1258 | 959 | 63 | 63 |  | 10575 | -9603 | -10 |  |  | A_LL4_G10_2 |  | 12.6896 | 29 | 2019-08-15 06:37
30 / 1 . | 571 | 265 | 365 | 460 | 413 | 169 | 841 | 443 | 1267 | 967 | 63 | 65 |  | 6185 | -31662 | -83 |  |  | B_LL7_G5_1 |  | 12.5778 | 30 | 2019-08-15 06:46
31 / 1 . | 637 | 266 | 324 | 463 | 406 | 166 | 800 | 423 | 1197 | 1021 | 62 | 66 |  | 6218 | -31662 | -83 |  |  | B_LL7_G5_2 |  | 12.5718 | 31 | 2019-08-15 06:54
32 / 1 . | 587 | 266 | 324 | 481 | 376 | 166 | 824 | 442 | 1326 | 867 | 61 | 64 |  | 8053 | -31849 | -84 |  |  | B_LL7_G6_1 |  | 12.6615 | 32 | 2019-08-15 07:03
33 / 1 . | 590 | 267 | 303 | 484 | 381 | 167 | 835 | 433 | 1285 | 887 | 62 | 67 |  | 8086 | -31849 | -84 |  |  | B_LL7_G6_2 |  | 12.6667 | 33 | 2019-08-15 07:11
34 / 1 . | 625 | 258 | 334 | 468 | 405 | 164 | 862 | 435 | 1324 | 978 | 61 | 65 |  | 11195 | -31735 | -99 |  |  | B_LL7_G7_1 |  | 12.4512 | 34 | 2019-08-15 07:19
35 / 1 . | 626 | 262 | 338 | 504 | 398 | 167 | 840 | 419 | 1167 | 1052 | 61 | 64 |  | 11229 | -31735 | -99 |  |  | B_LL7_G7_2 |  | 12.4836 | 35 | 2019-08-15 07:28
36 / 1 . | 578 | 263 | 347 | 492 | 380 | 168 | 773 | 414 | 1175 | 941 | 60 | 66 |  | 17745 | -32449 | -117 |  |  | B_LL7_G10_1 |  | 12.4577 | 36 | 2019-08-15 07:36
37 / 1 . | 583 | 264 | 347 | 475 | 413 | 166 | 809 | 407 | 1399 | 933 | 63 | 63 |  | 17745 | -32466 | -117 |  |  | B_LL7_G10_2 |  | 12.4913 | 37 | 2019-08-15 07:44
38 / 1 . | 590 | 268 | 312 | 463 | 378 | 167 | 808 | 436 | 1320 | 932 | 62 | 65 |  | 6314 | -35077 | -79 |  |  | B_LL7_G12_1 |  | 12.4992 | 38 | 2019-08-15 07:53
39 / 1 . | 628 | 268 | 320 | 476 | 359 | 172 | 817 | 427 | 1308 | 1055 | 63 | 64 |  | 6314 | -35091 | -79 |  |  | B_LL7_G12_2 |  | 12.5471 | 39 | 2019-08-15 08:01

## Detection_limits_Plag
Det.Lim ppm
DataSet/Point | Na | Si | Mg | Al | K | Ca | Ti | Fe | Mn | O | X | Y | Z |  | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 691 | 798 | 183 | 654 | 356 | 837 | 112 | 661 | 758 |  | 7328 | 22323 | -77 |  |  |  | Lab_hole6_xstal1 |  | 11.6376 | 1 | 2020-01-31 09:56
2 / 1 . | 621 | 760 | 174 | 667 | 340 | 864 | 113 | 703 | 708 |  | 7365 | 22322 | -77 |  |  |  | Lab_hole6_xstal1_2 |  | 11.6883 | 2 | 2020-01-31 10:01
3 / 1 . | 668 | 804 | 175 | 655 | 343 | 847 | 112 | 663 | 826 |  | 7377 | 22331 | -77 |  |  |  | Lab_hole6_xstal2 |  | 11.7136 | 3 | 2020-01-31 10:05
Det.Lim ppm
DataSet/Point | Na | Al | P | Si | Mg | K | Ca | Ti | Fe | Mn | S | Cl | O | X | Y | Z | Beam X | Beam Y | Comment | Distance (µ) | Mean Z | Point# | Date
1 / 1 . | 664 | 279 | 375 | 843 | 330 | 476 | 385 | 162 | 1725 | 873 | 62 | 66 |  | 13375 | 5753 | 6 |  |  | LL2_443 |  | 11.8511 | 1 | 2020-02-02 04:05
2 / 1 . | 626 | 281 | 324 | 866 | 370 | 488 | 390 | 163 | 1741 | 894 | 62 | 66 |  | 11806 | 5119 | 22 |  |  | LL2_444 |  | 12.3373 | 2 | 2020-02-02 04:14
3 / 1 . | 648 | 279 | 324 | 857 | 314 | 478 | 359 | 156 | 1830 | 796 | 63 | 65 |  | 10648 | 4811 | 16 |  |  | LL2_445 |  | 11.654 | 3 | 2020-02-02 04:22
4 / 1 . | 634 | 293 | 372 | 821 | 316 | 420 | 407 | 163 | 1702 | 881 | 63 | 64 |  | 12395 | 3572 | 22 |  |  | LL2_447_A |  | 12.0517 | 4 | 2020-02-02 04:31
5 / 1 . | 620 | 294 | 354 | 803 | 333 | 462 | 397 | 165 | 1924 | 960 | 62 | 67 |  | 12173 | 3530 | 24 |  |  | LL2_447_D |  | 11.902 | 5 | 2020-02-02 04:39
6 / 1 . | 607 | 280 | 338 | 838 | 338 | 508 | 409 | 163 | 1876 | 837 | 63 | 66 |  | 10703 | 3423 | 24 |  |  | LL2_448 |  | 12.0089 | 6 | 2020-02-02 04:47
7 / 1 . | 730 | 296 | 343 | 877 | 326 | 467 | 376 | 160 | 1190 | 775 | 63 | 65 |  | 8912 | 2940 | 4 |  |  | LL2_449 |  | 11.819 | 7 | 2020-02-02 04:56
8 / 1 . | 649 | 295 | 363 | 846 | 292 | 479 | 346 | 156 | 1809 | 934 | 62 | 66 |  | 9121 | 2455 | 4 |  |  | LL2_450_a |  | 11.5274 | 8 | 2020-02-02 05:04
9 / 1 . | 652 | 282 | 400 | 837 | 305 | 506 | 365 | 162 | 1517 | 937 | 62 | 66 |  | 9367 | 2564 | 4 |  |  | LL2_450_c |  | 11.6759 | 9 | 2020-02-02 05:12
10 / 1 . | 663 | 287 | 291 | 842 | 345 | 491 | 368 | 162 | 1859 | 858 | 62 | 66 |  | 9281 | 2695 | 4 |  |  | LL2_450_b |  | 11.6321 | 10 | 2020-02-02 05:21
11 / 1 . | 672 | 292 | 314 | 876 | 303 | 490 | 398 | 162 | 1787 | 948 | 60 | 66 |  | 16953 | 3150 | 24 |  |  | LL2_451_a |  | 11.7248 | 11 | 2020-02-02 05:29
12 / 1 . | 686 | 291 | 323 | 779 | 403 | 374 | 454 | 179 | 2054 | 892 | 61 | 64 |  | 14551 | 1531 | 11 |  |  | LL2_452 |  | 12.3253 | 12 | 2020-02-02 05:38
13 / 1 . | 648 | 284 | 332 | 835 | 348 | 513 | 404 | 163 | 1675 | 906 | 63 | 67 |  | 12393 | 1585 | 9 |  |  | LL2_453_b |  | 12.0347 | 13 | 2020-02-02 05:46
14 / 1 . | 592 | 281 | 348 | 846 | 343 | 488 | 358 | 166 | 1799 | 814 | 62 | 66 |  | 12278 | 1805 | 9 |  |  | LL2_453_a |  | 11.9924 | 14 | 2020-02-02 05:55
15 / 1 . | 664 | 288 | 365 | 884 | 323 | 507 | 394 | 163 | 1860 | 766 | 63 | 66 |  | 12445 | 1946 | 15 |  |  | LL2_453_c |  | 11.9697 | 15 | 2020-02-02 06:03
16 / 1 . | 617 | 294 | 399 | 808 | 403 | 370 | 452 | 178 | 1731 | 797 | 62 | 65 |  | 19172 | 2783 | 10 |  |  | LL9_467 |  | 12.6718 | 16 | 2020-02-02 06:11
17 / 1 . | 682 | 290 | 393 | 793 | 375 | 399 | 444 | 179 | 1983 | 830 | 65 | 66 |  | 15397 | 514 | 6 |  |  | LL9_474_c |  | 12.7871 | 17 | 2020-02-02 06:20
18 / 1 . | 656 | 296 | 377 | 794 | 396 | 408 | 436 | 174 | 1838 | 976 | 65 | 66 |  | 14664 | 540 | 11 |  |  | LL9_476 |  | 12.6139 | 18 | 2020-02-02 06:28
19 / 1 . | 645 | 280 | 304 | 784 | 412 | 421 | 455 | 182 | 1961 | 887 | 64 | 67 |  | 12614 | -134 | -2 |  |  | LL9_478 |  | 12.8144 | 19 | 2020-02-02 06:37
20 / 1 . | 615 | 289 | 374 | 799 | 375 | 409 | 448 | 180 | 2033 | 849 | 62 | 65 |  | 15737 | -829 | -4 |  |  | LL11_492 |  | 12.7385 | 20 | 2020-02-02 06:45
21 / 1 . | 564 | 286 | 388 | 765 | 411 | 407 | 441 | 175 | 1618 | 914 | 64 | 66 |  | 16365 | -1314 | -7 |  |  | LL11_493_c |  | 12.6676 | 21 | 2020-02-02 06:53
22 / 1 . | 617 | 297 | 375 | 775 | 415 | 379 | 458 | 179 | 1877 | 848 | 64 | 66 |  | 12871 | -1667 | -2 |  |  | LL11_495_a |  | 12.8489 | 22 | 2020-02-02 07:02
23 / 1 . | 688 | 298 | 400 | 795 | 397 | 417 | 457 | 176 | 2038 | 946 | 63 | 66 |  | 13094 | -1461 | . |  |  | LL11_495_b |  | 12.8388 | 23 | 2020-02-02 07:10
24 / 1 . | 647 | 293 | 345 | 745 | 413 | 389 | 437 | 180 | 1769 | 960 | 65 | 65 |  | 12954 | -1380 | . |  |  | LL11_495_d |  | 12.6352 | 24 | 2020-02-02 07:18
25 / 1 . | 665 | 299 | 343 | 780 | 400 | 408 | 455 | 177 | 2033 | 899 | 63 | 67 |  | 13132 | -1247 | -2 |  |  | LL11_495_c |  | 12.6274 | 25 | 2020-02-02 07:27
26 / 1 . | 629 | 296 | 335 | 795 | 394 | 383 | 473 | 177 | 1989 | 897 | 65 | 67 |  | 13288 | -1400 | -2 |  |  | LL11_495_e |  | 12.7882 | 26 | 2020-02-02 07:35
27 / 1 . | 678 | 291 | 367 | 825 | 386 | 403 | 458 | 181 | 2010 | 855 | 65 | 66 |  | 10207 | -1886 | -24 |  |  | LL11_499 |  | 12.909 | 27 | 2020-02-02 07:44
28 / 1 . | 676 | 295 | 375 | 782 | 430 | 423 | 469 | 171 | 1847 | 788 | 63 | 66 |  | 8629 | -1350 | -24 |  |  | LL11_500_b |  | 12.6365 | 28 | 2020-02-02 07:52
29 / 1 . | 622 | 302 | 386 | 791 | 404 | 411 | 438 | 173 | 1632 | 824 | 63 | 66 |  | 17700 | -2188 | -4 |  |  | LL12_510_b |  | 12.4112 | 29 | 2020-02-02 08:00
30 / 1 . | 556 | 296 | 361 | 825 | 370 | 432 | 428 | 177 | 1986 | 922 | 63 | 68 |  | 17031 | -2503 | -11 |  |  | LL12_513_c |  | 12.6149 | 30 | 2020-02-02 08:09
31 / 1 . | 622 | 295 | 361 | 813 | 383 | 414 | 424 | 169 | 1916 | 835 | 63 | 66 |  | 16868 | -2616 | -11 |  |  | LL12_513_b |  | 12.2975 | 31 | 2020-02-02 08:17
32 / 1 . | 653 | 293 | 308 | 785 | 396 | 449 | 457 | 176 | 1885 | 896 | 65 | 64 |  | 15152 | -3248 | -20 |  |  | LL12_514 |  | 12.6214 | 32 | 2020-02-02 08:26